## Slide 1: Production of metal ion beams from DECRIS-3 ion source
А.E. Bondarchenko1, S.L. Bogomolov1, V.N. Loginov1, A.N. Lebedev1, V.E. Mironov1, D.K. Pugachev1, М.V. Zdorovets2, I.А. Ivanov2 ,Е.К. Sambayev2, M. V. Koloberdin2,А.Е. Kurakhmedov2, D.A. Mustafin2, M.B. Abdigaliyev2
1 Joint Institute for Nuclear Research, FLNR, Dubna, Russia
2 Astana branch of Institute of Nuclear Physics, Nur-Sultan, Kazakhstan

## Slide 2: The accelerator complex DC-60 of Astana branch of the INP (Nur Sultan, Kazakhstan Republic)
The main parameters of the beams
Type of ions | Li - Xe
Mass to charge ratio (A/Z) | 6 - 12
Accelerated ion energy | 0,35 – 1,7 MeV/nucleon
MAIN OBJECTIVES :
Scientific research
Education
Production of track membranes with special properties-
Creation of micro and nano structures
Surface modification of standard materials, creation of new materials with required properties
The cyclotron is equipped with the Electron Cyclotron Resonance (ECR) ion source DECRIS-3

## Slide 3: CYCLOTRON DC-60
DECRIS-3 ion source
The experience and knowledge on the production of metal ion beams from ECR ion sources in FLNR by the MIVOC method has been successfully transferred to the DC-60 cyclotron.

## Slide 4
The ECR ion source

## Slide 5
Production of ions of solids
Many of the elements required for acceleration at the cyclotrons are available in the solid state form only.

## Slide 6
Vapor of metals
(oven method)
Volatile compounds
(MIVOC method)
Production of ions of solids
or
FLNR microoven
Metals vapor pressure vs temperature
Oven method

## Slide 7
Ion | Energy, MeV/nucleon | Ion beam current in injection, µA | The extracted ion beam current, µA | Transmission ratio, %
40Ca7+ | 1,75 | 50 | 1,35 | 2,75
31P5+ | 1,75 | 20 | 0,58 | 3,00
24Mg4+ | 1,75 | 46 | 1,14 | 2,48
7Li+ | 1,32 | 100 | 1,5 | 1,5
Parameters of accelerated ions: 40Ca7+, 24Mg4+, 31P5+, 7Li+
Metal ion beams produced at DC—60 cyclotron using oven method

## Slide 8
MIVOC-method (Metal Ions from VOlatile Compounds)
Requirements:
1) Chemical compound containing required element has to be found
2) The saturated vapour pressure (at room temperature) of the compound has to be high enough (≥ 10-3 mbar)
- minimum usable vapour pressure of the compound is determined by the material consumption rate of the ion source and by the conductance between the MIVOC chamber and the ion source
C2B10H12 - carboran
(C5H5)2Me – metallocene molecule (Fe, Ni,Cr..)
(CH3)4Me – tetrameethyl molecule (Sn, Ge)
(CO)6Me – hexacarbonyl molecule (Cr, W)
W(CO)6
Fe(C5H5)2

## Slide 9
Titanium ions beams were produced from the compound (CH3)5C5Ti(CH3)3.
Preparation of the titanium compound (CH3)5C5Ti(CH3)3 for operation. Left - an ampoule with a substance is placed in a container; in the center - the container with the substance is pumped to the forevacuum; on the right - the container with the working substance is connected to the ECR ion source through a mechanical regulating valve.
The preparation of the titanium compound.
The preparation compounds ferrocene, cobalt, chrome, nickel, germanium, silicium and hafnium.
The organometallic compound Fe(C5H5)2 hasn’t a chemical activity with oxygen and isn’t sensitive to visible light. Therefore, the preparation of the MIVOC chamber with the compound Fe(C5H5)2 was carried without an argon box.
The compounds Cr(C5H5)2, Co(C5H5)2, Ge(CH3)4, Si(CH3)4 and (C­5H5)2Hf(CH3)2 has a chemical activity with oxygen. So, the preparation of the MIVOC chamber was carried out an argon box.
The compound Ni(C5H5)2 has a high chemical activity with oxygen and is highly sensitive to visible light. Therefore, the preparation of the MIVOC chamber with the compound Ni(C5H5)2 was carried out in an argon box under red light.
After preparation all containers with compounds are pumped to the forevacuum, and connected to the ECR ion source through a mechanical regulating valve.

## Slide 10
The spectrum of Cr ions, the source settings are optimized for 52Cr10+ - 18 μA . UHF power – 72 W.
The spectrum of Hf ions, the source settings are optimized for 180Hf17+ - 19 μA. UHF power – 230 W.

## Slide 11
The spectrum of Si ions, the source settings are optimized for 28Si5+ - 62 μA. UHF power – 50 W.
The spectrum of Ti ions, the source settings are optimized for 48Ti9+ - 18 μA.
UHF power – 156 W.

## Slide 12
The spectrum of Co ions, the source settings are optimized for 59Co12+ - 12 μA .
UHF power – 207 W.
The spectrum of Ni ions, the source settings are optimized for 58Ni11+ - 43 μA. UHF power – 260 W.

## Slide 13
The spectrum of Ge ions, the source settings are optimized for 72Ge10+ - 46 μA. UHF power – 270 W.
The spectrum of Fe ions, the source settings are optimized for 56Fe10+ – 44 µA. UHF power – 120 W.

## Slide 14
Table. The intensity (μA) of metal ion beams produced from the DECRIS-3 ECR ion source of the DC-60 cyclotron using the MIVOC method
Z | 5+ | 6+ | 8+ | 9+ | 10+ | 11+ | 12+ | 14+ | 15+ | 17+ | 18+ | 20+
28Si | 62* | 40 |  |  |  |  |  |  |  |  |  | 
48Ti |  |  |  | 18* | 16 |  |  |  |  |  |  | 
58Ni |  |  |  | 39 |  | 43* | 37 |  |  |  |  | 
59Co |  |  |  | 57 | 73 | 36 | 12* |  |  |  |  | 
52Cr |  |  | 22 |  | 9* |  |  |  |  |  |  | 
72Ge |  |  | 38 |  | 46* | 34 |  |  |  |  |  | 
56Fe |  |  | 68 |  | 44* | 25 | 12 |  |  |  |  | 
180Hf |  |  |  |  |  |  |  |  |  | 19 | 15 | 5
* –metal ion beams accelerated at the DC-60 cyclotron

## Slide 15: CONCLUSION
The main goal of this work was to employ the MIVOC method for production of ions of solid and extend the spectrum of accelerated ions at the DC-60 cyclotron.
With the expansion of the spectrum of accelerated elements, it becomes possible to set up new experiments in the field of experimental nuclear physics, radiation solid state physics, and various applied problems, which makes the research objective particularly relevant. As a result of the work done, for the first time at the DC-60 cyclotron, the ion beams of 48Ti9+, 72Ge10+, 56Fe10+, 180Hf17+, 58Ni11+, 28Si5+, 59Co12+ and 52Cr10+ were produced.

## Slide 16: Thanks for attention