## Slide 1: Report on the infrastructure developments, including the Nuclotron
N.N.Agapov

## Slide 2: PRESENTATION OUTLINE
Main directions of the NICA infrastructure development:
Buildings
Line for assembling and cryogenic testing of SC-magnets
Water chilling systems on the base of open cooling towers
Cryogenics - liquid helium and liquid nitrogen plants
Modernization of electricity systems
Modernization of heating facilities
Reconstruction of the water pipelines, electric power networks and thermal grids

## Slide 3: Report on the construction of the collider building
December 2016:
- total number of organizations - 9
(general contractor of "Strabag" and 8 sub-contractors)
- total number of engineers and workers - 96

## Slide 4: Report on the construction of the collider building

## Slide 5
Existing compressor building 4,4 MW
New compressor building
8,6 MW

## Slide 6
Helium screw compressors “Kaskad-110/30”
Nitrogen turbo compressors “Aerocom2-179/18” 2 units in operation, 1 in reserve
Capacity of compressor, Nm3/h | 10740
Inlet pressure, MPa | 0.102
Inlet temperature, °C | 30
Outlet pressure, MPa | 1.8
Outlet temperature, °C | 40
Temperature of cooling water, °C | 20
Installed power of electric motor, kW | 1800
Capacity (Nm3/h) | 6600
Outlet pressure (MPa) | 3.0
Total power of electric motors (kW) | 1600
Voltage (V) | 6000
Number of compression stages | 2
Speed (rpm) | 2970
Flow rate of cooling water, m3/h | 78
Nitrogen receivers
1300 kg/h nitrogen liquefier
30 m3 LN2 reservoir with pumps
200 m3 nitrogen gas bags

## Slide 7
In the framework of the Agreement between
JINR and Russian Federation about Mega-project NICA establishment of the centre of innovative developments has been planned.
- office space for staff (~ 300 seats according to the norms);
- computer centre with an area of 500 - 700 sq. m.;
- 4 laboratory facilities to organize applied research on ~ 50 sq. m and height up to 5m;
- meeting rooms ~ 50 sq. m (each floor except the first);
- necessary infrastructure facilities (toilets, lifts …)
- congress hall - transformer for 800 - 1000 people with possibility to divide into 3 or 4 parts;
- buffet area ~ 120 sq. m.

## Slide 8: Innovation center of NICA
At this time, several architectural solutions proposed by different project organizations are actively discussed. Here are two examples of such solutions.
Examples
of architectural solutions

## Slide 9
New heating trunk pipelines
Over the past few years, more than
2 km of the heating trunk pipelines have been reconstructed.
Enclosed in a ring, they will guarantee a reliable heating of all existing lab’s buildings.

## Slide 10: Modernization of heating facilities
Existing lines
New lines
Modernization needed
D500>D600mm
The heat load of the new buildings will be 11.7 Gcal, ie total load at VBLHEP will be doubled since existing value.
Eastern Thermal Station
A
B

## Slide 11: Official ceremony of launching of the high-technology line for assembling and testing of SC magnets was held. The ceremony was attended by leaders of the projects NICA and FAIR. November, 28

## Slide 12: Main production areas
Incoming inspection zone
SC cable production hall
SC coils production hall
Area for assembling of magnet
Area for magnetic measurement under the room temperature
Bay for leakage test
Area for mounting SC-magnet inside cryostat
Cryogenic tests bench

## Slide 13: Production plans
40 dipole magnets for NICA Booster
48 quadrupole magnets with multipole correctors for NICA Booster
80 dipole magnets for NICA Collider
86 quadrupole magnets with multipole correctors for NICA Collider
175 quadrupole magnets with multipole correctors for SIS100 synchrotron (FAIR project)
Total 429

## Slide 14: Cool down processes of the magnets

## Slide 15
At the present time the commissioning of the basic elements of the circulating water supply system for the Nuclotron machine hall is carried out.
Water chilling systems on the base of open cooling towers

## Slide 16
PUMPS
NB100-250/229
Cooling towers VLT-E
Vessel 20 м3

## Slide 17
Open cooling towers VTL-E
20 m3 vessel

## Slide 18
Cryogenics of the NICA accelerator complex
satellite refrigerator
of the collider
40 m3 liquid helium tank
1300 kg/h nitrogen liquefier OA-1.3
draining and oil-purification units MO-800
1000 l/h helium liquefier OG-1000
2x6600 Nm3/h helium screw compressors
3x10740 Nm3/h nitrogen turbo compressors
500 kg/h nitrogen re-condenser RA-0.5 of the collider
500 kg/h nitrogen re-condenser RA-0.5 of the Nuclotron and booster
Nuclotron
Booster
Collider
satellite refrigerator
of the booster

## Slide 19: Successful first run of the 1000 l/h helium liquefier
The largest helium liquefier in Russia was installed and successfully launched at JINR

## Slide 20: Simulated results of three-minute peak heat load during injection into Collider
Time-averaged heat load – 6700 W
Peak heat load – 9700 W
Transition process during peak load

## Slide 21: Reconstruction plans for electric networks and substations: - cable lines - power transformer station (110/6 kV) - substation 13.

## Slide 22: Construction of two parallel 6 kVcable lines from the Main power transformer station (110/6 kV) to Substation 13.
Main power transformer station (110/6 kV)
Substation 13

## Slide 23
Modernization of the electricity systems
The existing power transformer station is 19 megawatts. This capacity is almost completely used.
Estimated increase in capacity:
For the booster 1,2 MW
For the collider 6,0 MW
New compressor station 8,6 MW
Total 15,8 MW
For more power, it was decided to replace three existing transformers by two more powerful. After that the allowable load will increase to the required value.
For more power, it was decided to replace three existing transformers by two more powerful. After that the allowable load will increase to the required value.
For more power, it was decided to replace three existing transformers by two more powerful. After that the allowable load will increase to the required value.

## Slide 24
JSC "Siemens Transformers" (Voronezh, Russia)
Rated power 40 MVA
Maximum rated voltage 126 kV
Total mass with oil 52 t
Mass of oil 11 t
Shipping weight 45 t
Dimensions 5.5x3.5x5.0 m

## Slide 25: I hope that was able to show that all directions of the infrastructure at the Lab are developing at the right pace

## Slide 26
Thank you for your attention!