## Slide 1: «Increasing Bandwidth of data acquisition systems on IBR-2 reactor spectrometers in FLNP»
Presenters: Shvetsov V. V.
Primary author: Drozdov V.A.
Co-author: Shvetsov V. V.
NEC’2017
E-mail: shvetc_vas@mail.ru

## Slide 2: The relevance of the tasks
Modern detector systems with point detector elements currently used on and being developed for FLNP spectrometers can be connected to data acquisition systems based on MPD-240 and De-Li-DAQ 2D modules, respectively. The MPD-240 module enables data acquisition from a maximum of 240 elements with a maximum load of up to 8 mega events per second. This requires the bandwidth of computer communication channels of up to 50 MB/s, which is impossible with the USB 2.0 interface.
The device under development will increase 10 times the capacity of data acquisition systems.
2
NEC’2017

## Slide 3: Goals and objectives
This work describes development of the new optical converter FLINK USB 3.0 to connect MPD-240 and De-Li-DAQ 2D modules with a computer through an USB 3.0 interface.
Variants of using USB 3.0 interface for the existing data acquisition systems in FLNP are considered;
A new FPGA project design of the optical converter is described;
The results of FPGA design debugging on the test development kit and on the first developed and manufactured samples of the converter are presented.
3
NEC’2017

## Slide 4
Neutrons from the reactor (1) go through the chopper (2) along the neutron guide (3) to the sample, irradiate the sample (4) and are scattered on it due to diffraction. The diffracted neutrons are registered by the detector system (5). Signals from the detector system, starts of the reactor and chopper come to the data registration system built on the basis of the MPD-240 module allowing “raw” data collection in the list mode. In such a system processing and computation of high resolution spectra are carried out off-line.
4
NEC’2017

## Slide 5: Types of data acquisition systems on the IBR-2 spectrometers
5
MPD-16
NEC’2017

## Slide 6: MPD-240
6
NEC’2017

## Slide 7: De-Li-DAQ 2D
7
NEC’2017

## Slide 8: FLINK
8
NEC’2017

## Slide 9
9
The data from the detectors are transmitted to the data acquisition system, then with a speed of 1.25 Gb/s to the FLINK module by the optical interface and finally, with a speed of 10 MB/s to the computer memory by the serial USB 2.0 interface.
DAQ
F-link
PC
D1
Dn
USB 2.0
Fiber-optic cable
1.25 Gibit/s
10 MB/s
NEC’2017

## Slide 10: IBR-2 facilities
10
NEC’2017

## Slide 11: USB 3.0. advantages
11
Maximum effective data rate is 375 MB/s
Malfunction or error probability is lower than with the USB 2.0 interface due to additional twisted pair shielding;
Feed current increases from 500 mA to 900 mA.
NEC’2017
Within the framework of the problem of USB 3.0 integration, the following options are considered:

## Slide 12: Attaining the objectives
12
Variant 1:
NEC’2017
DAQ
F-link
PC
D1
Dn
USB 2.0
Fiber-optic cable
1.25 Gibit/s
10 MB/s
375 MB/s
Fiber-optic to USB 3.0 converters
USB 3.0
USB 3.0
Fiber-optic cable

## Slide 13: USB 3.0 to Fiber Optical extender
13
NEC’2017

## Slide 14
DAQ
PC
D1
Dn
Fiber-optic cable
375 MB/s
Fiber-optic to USB 3.0 converters
USB 3.0
USB 3.0
Fiber-optic cable
14
Development of new blocks with USB 3.0 interface is more promising. This permits increasing the data transmission bandwidth up to 375 MB/s preserving the possibility of remote connection as required for FLNP spectrometers
In developing the new blocks there will be taken into account new requirements for characteristics of the existing data acquisition systems that arose in the process of their operation as well as to meet the requirements of the newly developed projects.

## Slide 15: Attaining the objectives
15
Variant 2:
DAQ
F-link
PC
D1
Dn
USB 2.0
Fiber-optic cable
1.25 Gibit/s
10 MB/s
USB 3.0
100 MB/s
NEC’2017
2.5 Gibit/s
This will allow us to use a maximum data transmission speed of the fiber interface of old blocks and will increase 10 times the bandwidth of the data transmission channel.

## Slide 16: Advantages and disadvantages
16
advantages | disadvantages
Easy commissioning | Transmission speed is lower than with direct connection
variant 2
advantages | disadvantages
Transmission speed up to 375 MB/s | Complexity of commissioning
Simplified architecture | 
variant1
NEC’2017
The second option allows increasing the data transmission rate by developing just one module, which allows more effective use of all the modules already installed on the spectrometers.

## Slide 17: Solution
17
To date, data transmission at 100 MB per second satisfies the requirements for the volume of the transmitted data. Moreover, integration of new FLINK modules will go faster. So, such work was the first to be performed.
NEC’2017

## Slide 18: FLINK USB 3.0 architecture
18
NEC’2017
In the framework of solving this problem the architecture of the new module consisting of three main integrated has been developed:

## Slide 19: FLINK USB 3.0
19
NEC’2017
To date, first five samples of FLINK-USB 3.0 converters have been manufactured.

## Slide 20: Test kit
20
NEC’2017
The FT600/FT601 is a USB 3.0 to FIFO interface bridge chip with the following advanced features: Supports USB 3.0 Super Speed (5Gbps)/USB 2.0 High Speed (480Mbps)/USB 2.0 Full Speed (12Mbps) transfer. Supported USB Transfer Type: Control/Bulk/Interrupt Up to 8 configurable endpoints (PIPEs). Supports 2 parallel slave FIFO bus protocols 245 and FIFO mode, FT601 with 32 bit parallel interface has a data bursting rate up to 400MB/s. Supports 4 IN channels and 4 OUT channels on FIFO bus connectivity. Built-in 16kB FIFO data buffer RAM.
FTDI UMFT 600A
Cyclone V debugging set
Optical interface module

## Slide 21: Design of firmware FLINK USB 3.0
21
NEC’2017

## Slide 22
Operation in LoopBack mode
22
NEC’2017
This speed is a maximum feasible one in transmitting data through TLK1501.

## Slide 23: FLINK USB 3.0
23
NEC’2017
During the first cycle of reactor operation the system is to be tested on the spectrometers of the IBR-2 reactor in real experimental conditions.

## Slide 24: Results
24
As a result, the variant of increasing by developing a new optical convertor, FLINK USB 3.0, has been realized.
  High-speed chip from FTDI FT600 is used;
  FPGA Cyclone IV EP4CE6E22C8N is used;
  New FPGA project design for data exchange between the computer and the data acquisition system via the USB 3.0 interface is developed.
The testing kit based on Cyclone V 5CGXF5C6F27C7N Starter Kit, debugging kit USB 3.0 FTDI UMFT600A and the Fiber Optic Transceiver kit is developed.
The FPGA project design with a new FLINK USB 3.0 module allows over 10 times increase of the bandwidth of the data transmission channel) from the data acquisition and processing systems De-Li-DAQ-2D and MPD.
NEC’2017

## Slide 25: MPD-16 USB 3.0
25
NEC’2017
Work to increase the bandwidth of data acquisition systems will continue. A new MPD-16 module with USB 3.0 interface has already been designed. Work on developing the firmware for MPD-240 and De-Li-DAQ modules is under way today.

## Slide 26
Thanks for your attention!
26

## Slide 27
NEC’2017
«Increasing throughput capacity of data acquisition systems on IBR-2 reactor spectrometers in FLNP»
Presenters: Shvetsov V. V.
Primary author: Drozdov V.A.
Co-author: Shvetsov V. V.
E-mail: shvetc_vas@mail.ru