## Slide 1: Beam diagnostics at the NICA
Evgeny Gorbachev, Valery Volkov, LHEP, JINR

## Slide 2: Content
Overview of diagnostic devices.
Diagnostics devices quantity and placement on accelerators and transfer channels.
Conclusions.

## Slide 3: Measured properties
Position – beam position monitors.
Current – transformers.
Profiles – profile meters, ionisation profile monitor.
Beam losses – neutron detector.
Luminosity – detector.

## Slide 4: Beam position monitor
Prototype of the elliptical BPM.

## Slide 5: Beam position monitor
Chart: 
Sum | 700 | 688 | 614 | 582 | 562 | 552 | 540 | 532 | 528 | 528 | 524 | 524 | 522 | 522 | 520 | 522 | 522 | 522 | 524 | 524 | 526 | 528 | 532 | 534 | 538 | 544 | 548 | 560 | 572 | 594 | 630 | 700 | 928
Delta | 344 | 332 | 278 | 242 | 214 | 192 | 168 | 148 | 128 | 108 | 92 | 72 | 54 | 38 | 20 | 6 | -14 | -34 | -52 | -68 | -86 | -104 | -120 | -138 | -158 | -176 | -196 | -216 | -240 | -266 | -302 | -360 | -504
Chart: 
Normalized | 0.49143 | 0.48256 | 0.45277 | 0.41581 | 0.38078 | 0.34783 | 0.31111 | 0.2782 | 0.24242 | 0.20455 | 0.17557 | 0.1374 | 0.10345 | 0.072797 | 0.038462 | 0.011494 | -0.02682 | -0.065134 | -0.099237 | -0.12977 | -0.1635 | -0.19697 | -0.22556 | -0.25843 | -0.29368 | -0.32353 | -0.35766 | -0.38571 | -0.41958 | -0.44781 | -0.47937 | -0.51429 | -0.5431
BPM prototype was measured with current pulses through movable wire and rods of different diameters.
Measurements show good linearity of the BPM prototype despite mechanical problems.
wire position
wire position
V, mV
∆/Σ

## Slide 6: Beam position monitor (Booster)
Improved design of Booster BPM – separate frame with mounted plates. Drawing of BPM in lenses module:

## Slide 7: Beam position monitor (Booster)
Frame assembly welded into the vacuum chamber. Preproduction unit is being manufactured in Bulgaria.

## Slide 8: Beam position monitor (Booster)
BPM unit between magnets:

## Slide 9: Beam position monitor (collider)
BPM for arcs have similar design. Straight sections BPMs are cylindrical with 90mm diameter.

## Slide 10: BPM signals processing
14 bit ADC
100MHz clock
Low noise input amplifiers
2 Gbps data throughput
Nuclotron Orbita system:

## Slide 11: BPM signals processing
Libera Hadron system:
Specifications:
16 input channels (4 per module)
16 bit ADC
250MS/s
6.5Gbps Rocket I/O
4GB memory per module
Data paths:
Broadband data (ADC@250MHz) -270MS (>1s of data), 4 channels
Bunch-by-bunch data – 200MS (>66s of data @1MHz bunch rep rate)
Slow data stream (10 S/s)
Fast data stream (10kS/s)
Output data:
Raw(A,B,C,D)
calculated beam position (X,Y)
Charge(SUM)
FFT, FFT peak

## Slide 12: Beam current measurements
New Parametric current transformer (NPCT, Bergoz) – DC non-destructive beam current measurement.
Full scale ranges | ±20mA, ±200mA, ±2A, ±20A
Bandwidth | DC to 10 kHz
Output | ±10V
Resolution | <5µArms/√Hz
Linearity error | <0.1%
Output accuracy | ±0.1% ± zero offset

## Slide 13: Beam Current measurements
DC current measurements by NPCT installed on Nuclotron.

## Slide 14: Beam Current measurements
Fast current transformer (FCT, Bergoz) – non-destructive instrument to observe longitudinal beam profiles.
Sensitivity | From 0.5V/A to 20V/A
Bandwidth | Upper cutoff frequency up to 2GHz
Droop | <6 %/us
Rise time | < 200 ps
Output | 50Ω terminated SMA jack connector

## Slide 15: Beam Current measurements
Beam current measurements by FCTs on Nuclotron and LU20-Nuclotron transfer channel.

## Slide 16: Beam profile measurements (channels)
Two planes.
32×32 wires on high vacuum ceramics frames.
Gold plated tungsten 50 um diameter.
2 mm between wires.
70 mm internal diameter.
ConFlat DN100 flange.
Electrical actuator.

## Slide 17: Beam profile measurements (channels)
First item was manufactured in Bulgaria and delivered to JINR.

## Slide 18: Beam profile measurements (channels)
Digital synchronization, data acquisition and processing system prototype:
NI PXI-1031 4-slot crate
NI PXI-8820 controller
NI PXI-6255 acquisition board
NI PXI-7952R FlexRIO board

## Slide 19: Beam profile measurements (rings)
Ionization Profile Monitor:
Detection of residual gas ionization using micro-channel plate (MCP) detector.
Up to 1e9 single charged beam intensity
0.5 mm spatial resolution

## Slide 20: Beam loss monitors
BDKN-96 probe: measurements of dose rate, power flux density and neutron dose.
Dose rate | 0.1 μSv·h-1 ÷ 0.1 Sv·h-1
Dose range | 0.1 μSv ÷ 1.0 Sv
Power flux density | 1 ÷ 104s-1·сm-2
Energy range of neutrons | 0.025 eV ÷ 14 MeV
Protection index | IP65
Overall dimensions, weight: | 295x142x100mm2.25kg
Damage prevention: vital for superconducting accelerators.
Diagnostics: position and time of losses

## Slide 21: Tune measurements
Kicker Z
Kicker X
Impedance
transformers
RF Amplifier AR 800A3A
NI PXIe-1085 crate ( PXIe system )
with modules:
FlexRIO Digitizer (developed at JINR)
NI PXIe-7975R FlexRIO module
PXI-6733 High-Speed Analog Outputs
NI PXIe-8135 2.3 GHz Quad-Core PXI Express Controller
FlexRIO
Digitizer
DAC 14 bit
output
Pick-up electrodes
X and Z
Amplifiers,
Filters,
Detectors
FlexRIO
Digitizer
ADC 18 bit
input

## Slide 22: HILAc-Booster channel diagnostics
Dipole magnet
Quadrupole
Debuncher
Dipole corrector
Faraday cup
CT
Pick-up
PU in corrector
Profile meter
Profile+Faraday cup
Collimator
Valve
BPM(X/Z) | 4
Current transformer | 3
Faraday cup | 4
Beam profile monitor | 5

## Slide 23: Booster-Nuclotron channel diagnostics
Under design in Novosibirsk.
No information about beam diagnostics yet.

## Slide 24: Nuclotron-collider channel diagnostics
The transfer channel will be manufactured by Sigmaphi.
BPM(X/Z) (Libera Spark HL) | 15
Beam loss monitor (Libera BLM) | 10
Fast current transformer | 4
Beam profile monitor (Libera Photon) | 10

## Slide 25: Booster diagnostics layout
D
F
D
F
D
F
D
F
D
F
D
M
M
M
M
M
M
M
M
M
M
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpm
bpm
bpm
bpm
bpm
Vacuum station
F quadrupole
D quadrupole
Magnet
Multipole corrector
BPM
scraper
One quadrant:
Lenses block:
BPM in lenses modules (X/Z) | 24
BPM in magnet modules (X/Z) | 20
Parametric current transformer (Bergoz) | 1
Fast current transformer (Bergoz) | 1
Betatron tune measurements (Q-meter) | 1
Ionization profile monitor | 1
Beam loss monitor | 24
F
Total number of diagnostic devices:
PU(Z)
PU(X)
D
F

## Slide 26: Collider diagnostics (arcs)
Collider cell:
MC- multipole corrector
CP - dipole corrector + PU

## Slide 27: Collider diagnostics (straight section)
MCP – multi corrector pack

## Slide 28: Collider diagnostics
Pick-up near F lenses (X) | 24
Pick-up near D lenses (Z) | 24
Pick-up in straight sections (X/Z) | 20
Parametric current transformer (Bergoz) | 1
Fast current transformer (Bergoz) | 2
Betatron tune measurements (Q-meter) | 1
Ionization profile monitor | 1
Beam loss monitors (for both rings) | 68
Total number of diagnostic devices in one ring (2 arcs + 2 straight sections):

## Slide 29: Conclusions
Overview of the basic diagnostic systems in Booster, Nuclotron, collider rings and transfer channels, their quantity and placement has been presented.
Thank you for your attention.