## Slide 1: Status of the Factory of Super Heavy Elements DC-280 cyclotron
Igor Kalagin
Dubna 2019
FLEROV LABORATORY of NUCLEAR REACTIONS
JOINT INSTITUTE FOR NUCLEAR RESEARCH

## Slide 2: Superheavy Elements (SHE) Factory – the Goals
Experiments at the extremely low (σ<100 fb) cross sections:
Synthesis of new SHE in reactions with 50Ti, 54Cr ...;
Synthesis of new isotopes of SHE;
Study of decay properties of SHE;
Experiments requiring high statistics:
Nuclear spectroscopy of SHE;
Study of chemical properties of SHE.

## Slide 3: SHE Factory
High-current cyclotron DC-280
New facilities:
New gas-filled separator
Preseparator
SHELS
Etc.
SHE Factory Building
Notes: The project of the SHE factory includes three principal components: i) high current accelerator, ii) new experimental hall and iii) a number of the next generation facilities.

## Slide 4
Stand-alone SHE factory with DC-280 cyclotron
DC280 (expected) E=4÷8 MeV/A |  | 
Ion | Ion energy [MeV/A] | Output intensity
7Li | 4 | 1×1014
18O | 8 | 1×1014
40Ar | 6 | 6×1013
48Ca | 6 | 6,2×1013
50Ti | 6 | 3,1×1013
54Cr | 6 | 2×1013
58Fe | 5 | 1×1013
124Sn | 5 | 2×1012
136Xe | 5 | 1×1014
238U | 7 | 5×1010
SHE factory building 2018

## Slide 5: Approval of the Statement of Compliance
Federal Service
For Environmental, Industrial and Nuclear Supervision
(Rostechnadzor)
Central Department
ORDER
20 November 2018 No. 386
Moscow
On approval of the statement of compliance of the built and reconstructed unit of capital construction with engineering requirements, other regulations and design documentation, including energy efficiency requirements and the requirements for the electricity meter infrastructure at the unit of capital construction
The statement of compliance as of 20 November 2018 No. 5.4-2943-пр-Зк/0117-2018 of the build and reconstructed unit of capital construction “Experimental building of the Laboratory of Nuclear Reactions”, located in Moscow region, 141980 Dubna, Joliot-Curie Str. 20, with engineering requirements, other regulations and design documentation, including energy efficiency requirements and the requirements for the electricity meter infrastructure at the unit of capital construction to be approved.
Customer: The Joint Institute for Nuclear Research (JINR, Primary State Registration Number PSRN 1035002200221, Taxpayer ID 9909125356, address: 141980, Moscow region, Dubna, Joliot-Curie Str. 6.

## Slide 6
Program of Complex Launching and Tuning Works (LTW) with ion beams on DC-280 agreed with the Federal Medical and Biological Agency of Russia (FMBA)
FMBA letter to JINR Director about possibility to start the Complex LTW with ion beam on DC-280 according with the agreed program

## Slide 7
Ion sources | Permanent magnet ECR DECRIS-PM - 14 GHz
Injection energy | Up to 80 keV/Z
A/Z range | 4÷7.5
Energy | 4÷8 MeV/n
Magnetic field level | 0.6÷1.3 T
K factor | 280
Magnet weight | 1000 t
Magnet power | 300 kW
Dee voltage | 2x130 kV
RF power consumption | 2x30 kW
Flat-top dee voltage | 2x14 kV
Deflector voltage | Up to 90 kV
DC-280
Main Parameters

## Slide 8
Main RF-resonator
Magnet yoke
HV platform N1
(Umax=70 kV)
Electrostatic deflector
(Bender)
Separating magnet
DECRIS-PM
ECR-source
(Umax=20 kV)
Magnet coils
Ion beam extraction
Flat –top resonator
Vacuum pump
Accelerating tube
Configuration of the DC-280

## Slide 9
Working diagram of the DC-280
First experiments at SHE factory

## Slide 10: Tests of DECRIS-PM at the HV platform of DC-280
Kr
Ar
84Kr+14 , A/Z=6
I max=65 µA
40Ar+7 , A/Z=5.71
I max=190 µA
Ions for DC-280 tests
40Ar+8, A/Z=5
I max=290 µA
-The first beam
UECR=15 kV

## Slide 11
Calculated regime of 84Kr+14 acceleration
ION | A/Z | Rextr | Rextr 120 kV | F | HARM | В0 | В at Rextr | Im | Uinj | Ubender | Uinfl | Udefl | E
 |  | m | mm | MHz |  | T | G | A | kV |  kV |  kV | kV | MeV/u
48Ca8+, 84Kr14+ | 6 | 1.76 | 12 | 9 | 3 | 1.1722 | 76.8 | 857 | 62.14 | 9.32 | 12.43 | 76 | 5.8
Gs
Lower activation of components at 6 Mev/u, good for materials outgazing, relatively low max ion current
The first beam: 84Kr+14

## Slide 12: Beam injection system
The HV platform (December 2018)
UDECRIS-PM=15 kV
UHV platform=47.5 kV
Ion beam: 84Kr+14
ILF1
ILF2
ILF3
Faraday cup | I (µA) | I (µA) | I (µA)
IFC1 | 40.7 | 30.1 | 18.1
IFC1 | 41 | 32.7 | 18.6
IFC1 | 40.8 | 30.8 | 17.7
HV
We used commissioning variant of inflector
regular inflector

## Slide 13
ILF1
ILF2
ILF3
~40
~40
Collimated
65 µA -> 41.8 µA
~40
~30
75
Transport of 84Kr+14 beam through injection system

## Slide 14
RF system
3Frf
Frf
Frf=7.32÷10.38 MHz
Fft=21.96÷31.14 MHz
Spectrums of X-rays from dee N1 for different dee voltages measured by XR-100T-CdTe detector
Dependence of maximal energy of X-rays on pick-up voltage
`

## Slide 15
26.12.2018
The first accelerated beam
of 84Kr+14 IRP1=1nA
R probe 1=281 мм
Im=831.9 A 10:31 a.m.
`
Current at radius of 1700 mm was up to 4 µA at injection current of 31 µA
Beam was not extracted due to problems with deflector
The first ion beam accelerated in DC-280

## Slide 16
Beam extraction system
Electrostatic deflector
Focusing magnetic channel
Extraction radius of Rext=1.79 m
Electrostatic deflector
L=1.3 m, Emax=90 кV/сm
Deflector tuning
Udefl=75.55 kV, Ileak=1.13 mA
`

## Slide 17
Radial trimming coils
Azimuthal trimming coils
Number of radial trimming coils | 11
Number of azimuthal trimming coils | 8
Magnet trimming coils

## Slide 18
B (T)
Resonance characteristics of 84Kr+14 acceleration
(Garren-Smith method)
50 G

## Slide 19
Transmission of 84Kr+14 at acceleration
Without buncher
With 1-st harmonic of buncher
Preliminary results
I FC1 | 18.1 | µA |  | Polyharmonic buncher | 
I FC2 | 18.6 | µA |  |  | 
I FC3 | 17.7 | µA |  |  | 
 | R, mm | I, µA | I/I FC1 |  | 
PR1 | 1400 | 2.8 | 16% | without buncher | 
PR1 | 1400 | 9.8 | 54% | with 1-st harmonic | 545 V
PR1 | 1400 | 10.3 | 57% | with 1-st and 2-nd harmonics | 900 V, 404 V
Capture: ~13%
Losses: ~40%
Capture: ~51%
Losses: ~20%

## Slide 20
Dependence of ion current on dee voltage

## Slide 21
17.01.2019
The first extracted beam
of 84Kr+14 IRP1=1nA
18:32
`
The first ion beam extracted from DC-280
without flat-top resonators

## Slide 22
Extraction channels
Bending magnet ±50
Quadrupole lenses
~45
~10
Beam of 100 nA on the TLF0 luminophor
T0LF1

## Slide 23
Vacuum in the DC-280 chamber PCG1=2.8610-7 Torr, PCG3=3.3810-7 Torr
`
Vacuum in the extraction channel: PT0G1=3.1610-7 Torr
Vacuum conditions
`

## Slide 24
Efficiency of 84Kr+14 beam transportation
from ECR source to extraction channel
Extraction without flat-top resonators
Measuring point | Beam intensity A | Transmission factor |  |  | 
ECR source, after separation | 10.6 | 17.6% |  |  | Total 6%
Cyclotron center | 1.87 |  | 48.7% |  | 
Extraction radius | 0. 91 |  |  | 70.3% | 
Extracted beam | 0.64 |  |  |  | 
Measuring point | Beam intensity A | Transmission factor |  |  | 
ECR source, after separation | 11 | 51.4% |  |  | Total 18.2%÷ 27.3%
Cyclotron center | 5.65 |  | 78.8% |  | 
Extraction radius | 4.45 |  |  | 45%÷ 67.4% | 
Extracted beam | 2÷3 |  |  |  | 
Transmission without buncher
Transmission with only the 1-st harmonic of buncher
Preliminary results
Notes: transmission factor.
Less than 20% of beam is lost in the process of acceleration
The overall efficiency is slightly less than 10 percent

## Slide 25
Radiation conditions
`
`
3.63 µZv/hour at 1 µA of extracted 84Kr+14 beam

## Slide 26
Future plans
Continuation of Complex LTW with acceleration of 12C+2, 40Ar+7, 48Ca+8,+9, 84Kr+14 ion beams. Carrying out radiation measurements with participation of FMBA representatives.
Installation of flat-top resonators, installation of regular inflector instead of commissioning variant, improving vacuum in the cyclotron chamber by searching hidden leak in the RF resonator N1.
Acceleration of 48Ca+8,+9 50Ti+8,+9. Increasing ion beam intensity to designed, transportation of ion beams to the GNS-2 experimental setup, measuring of beam transmission, measuring of beam parameters (energy, intensity distribution in cross-sections), preparation to first experiments.

## Slide 27
The first beam of 84Kr+14 ions was accelerated in the DC-280 cyclotron on 26.12.2018.
The first beam of 84Kr+14 ions was extracted from the DC-280 cyclotron on 17.01.2019.
Complex LTW with acceleration of 12C+2, 40Ar+7, 48Ca+8,+9, 84Kr+14 ion beams will be carried out up to March 2019.
The DC-280 will be prepared to first experiments with acceleration of 48Ca+8,+9 50Ti+8 ,+9 in March 2019.
Conclusion

## Slide 28
S.Dmitriev, Yu.Oganessian, G.Gulbekian, I.Kalagin, B.Gikal, S.Bogomolov, A.Efremov, A.Lebedev, V.Semin, V.Buzmakov, N. Pchelkin, S. Prokhorov, V. Sokolov, V.Mironov, N.Kazarinov, I.Ivanenko, V.Kazacha, V. Melnikov, O.Borisov, A.Tikhomirov, K. Gikal, K.Papenkov, K.Verlamov, A. Issatov, S.Mitrofanov, Yu. Teterev, N.Osipov, G. Ivanov, V.Bekhterev, V.Zarubin, V.Galinsky, S.Pashenko, S.Gusarov, V.Aleinikov, V.Verevochkin

## Slide 29
The first extraсted beam 17.01.2019
The first accelerated beam 12.12.2018
DC280 accelerator hall, PAC-40 26.06.2014
Magnet of DC280 18.01.2017
THANKS FOR YOUR ATTENTION!