## Slide 1: Beam diagnostics at the NICA – progress and tasks
Evgeny Gorbachev on behalf of team LHEP, JINR

## Slide 2: Content
Overview of diagnostic devices.
Diagnostics devices quantity and placement on accelerators and transfer channels.
Conclusions.

## Slide 3: Measured properties
Beam position – beam position monitors.
Beam current – transformers.
Beam profiles – profile meters, ionization profile monitors.
Beam losses – neutron detectors.
Luminosity – detector.

## Slide 4: Beam position monitor
Prototype of the elliptical BPM.

## Slide 5: BPM signals processing
Libera Hadron system:
Specifications:
16 input channels (4 per module)
16 bit ADC
250MS/s
6.5Gbps Rocket I/O
4GB memory per module
Data paths:
Broadband data (ADC@250MHz) -270MS (>1s of data), 4 channels
Bunch-by-bunch data – 200MS (>66s of data @1MHz bunch rep rate)
Slow data stream (10 S/s)
Fast data stream (10kS/s)
Output data:
Raw(A,B,C,D)
calculated beam position (X,Y)
Charge(SUM)
FFT, FFT peak

## Slide 6: BPM signals processing
Libera Hadron
• ICB – COM Express with
Intel processor
• Up to four BPMs
• Timing module (EVRX)
• GDX module
• Optional slots

## Slide 7: BPM signals processing
The external preamplifier mounted close to the BPM is in charged to provide high intensity and low noise signals to the Libera Hadron unit. A huge 110 dB beam dynamic can be coupled with the Amplifier 110.
Supported gain ranges: 60, 50, 40, 30, 20, 10, 0, -10, -20, -30, -40 and -50 dB
Gain error between channels: ±0.1 dB
Gain switching time: <1 us

## Slide 8: BPM signals processing
Bunch-by-bunch (BBB) algorithm:
Bunch detection.
Processing window calculation.
Baseline restorer (shoe-box).
Energy calculation.
Position calculation.
Output data rate = Bunch-by-bunch (varying)

## Slide 9: BPM signals processing
Timed BBB algorithm:
Bunch position synchronized with external RF clock.
RF clock can be delayed
Signal energy calculated in the gated region.
Narrow band analysis (low charges):
SDR receiver.
I & Q calculation.
Position calculation.
Output data rate = 50 kHz (fixed).

## Slide 10
BPM signals processing
Different impedance matching schemes:
no matching (signal distorted).
buffer amplifier (S/N is affected).
matching transformer (signal amplitude is affected).
Matching with 6:1 transformer

## Slide 11: BPM test bench
Tasks: find out the best sensitivity (lowest possible beam intensity) for 3 different algorithms of DSP and 3 different impedance matching schemes.

## Slide 12: Beam position monitor
Chart: 
Sum | 700 | 688 | 614 | 582 | 562 | 552 | 540 | 532 | 528 | 528 | 524 | 524 | 522 | 522 | 520 | 522 | 522 | 522 | 524 | 524 | 526 | 528 | 532 | 534 | 538 | 544 | 548 | 560 | 572 | 594 | 630 | 700 | 928
Delta | 344 | 332 | 278 | 242 | 214 | 192 | 168 | 148 | 128 | 108 | 92 | 72 | 54 | 38 | 20 | 6 | -14 | -34 | -52 | -68 | -86 | -104 | -120 | -138 | -158 | -176 | -196 | -216 | -240 | -266 | -302 | -360 | -504
Chart: 
Normalized | 0.49143 | 0.48256 | 0.45277 | 0.41581 | 0.38078 | 0.34783 | 0.31111 | 0.2782 | 0.24242 | 0.20455 | 0.17557 | 0.1374 | 0.10345 | 0.072797 | 0.038462 | 0.011494 | -0.02682 | -0.065134 | -0.099237 | -0.12977 | -0.1635 | -0.19697 | -0.22556 | -0.25843 | -0.29368 | -0.32353 | -0.35766 | -0.38571 | -0.41958 | -0.44781 | -0.47937 | -0.51429 | -0.5431
BPM prototype was measured with current pulses through movable wire and rods of different diameters.
Measurements show good linearity of the BPM prototype despite mechanical problems.
wire position
wire position
V, mV
∆/Σ

## Slide 13: Beam Current measurements
Fast current transformer (FCT, Bergoz) – non-destructive instrument to observe longitudinal beam profiles.
Sensitivity | From 0.5V/A to 20V/A
Bandwidth | Upper cutoff frequency up to 2GHz
Droop | <6 %/us
Rise time | < 200 ps
Output | 50Ω terminated SMA jack connector
* Graphs by R.Jones

## Slide 14: Beam Current measurements
Beam current measurements by FCTs on Nuclotron and LU20-Nuclotron transfer channel.

## Slide 15: Beam current measurements
New Parametric current transformer (NPCT, Bergoz) – DC non-destructive beam current measurement.
Full scale ranges | ±20mA, ±200mA, ±2A, ±20A
Bandwidth | DC to 10 kHz
Output | ±10V
Resolution | <5µArms/√Hz
Linearity error | <0.1%
Output accuracy | ±0.1% ± zero offset

## Slide 16: Beam Current measurements
DC current measurements by NPCT installed on Nuclotron.

## Slide 17: Beam profile measurements (channels)
Two planes.
32×32 wires on high vacuum ceramics frames.
Gold plated tungsten 50 um diameter.
2 mm between wires.
70 mm internal diameter.
ConFlat DN100 flange.
Electrical actuator.

## Slide 18: Beam profile measurements (channels)
Digital synchronization, data acquisition and processing system prototype:
NI PXI-1031 4-slot crate
NI PXI-8820 controller
NI PXI-6255 acquisition board
NI PXI-7952R FlexRIO board

## Slide 19: Beam profile measurements
Ionization Profile Monitor:
Detection of residual gas ionization using micro-channel plate (MCP) detector.
Up to 1e9 single charged beam intensity
0.5 mm spatial resolution
78Kr+26 beam acceleration (2,3 GeV/u), Nuclotron

## Slide 20: Beam loss monitors
BDKN-96 probe: measurements of dose rate, power flux density and neutron dose.
Dose rate | 0.1 μSv·h-1 ÷ 0.1 Sv·h-1
Dose range | 0.1 μSv ÷ 1.0 Sv
Power flux density | 1 ÷ 104s-1·сm-2
Energy range of neutrons | 0.025 eV ÷ 14 MeV
Protection index | IP65
Overall dimensions, weight: | 295x142x100mm2.25kg
Damage prevention: vital for superconducting accelerators.
Diagnostics: position and time of losses

## Slide 21: Tune measurements
Kicker Z
Kicker X
Impedance
transformers
RF Amplifier AR 800A3A
NI PXIe-1085 crate ( PXIe system )
with modules:
FlexRIO Digitizer (developed at JINR)
NI PXIe-7975R FlexRIO module
PXI-6733 High-Speed Analog Outputs
NI PXIe-8135 2.3 GHz Quad-Core PXI Express Controller
FlexRIO
Digitizer
DAC 14 bit
output
Pick-up electrodes
X and Z
Amplifiers,
Filters,
Detectors
FlexRIO
Digitizer
ADC 18 bit
input

## Slide 22: HILAc-Booster channel diagnostics
Dipole magnet
Quadrupole
Debuncher
Steerer
Faraday cup
Fast CT
Pick-up
Profile viewer
Collimator
Valve
BPM(X/Z) | 5
Current transformer | 3
Faraday cup | 3
Beam profile monitor | 4
Pump

## Slide 23: Booster-Nuclotron channel diagnostics
BPM(X/Z) | 6
Current transformer | 2
Beam profile monitor | 2
The transfer channel will be manufactured by Novosibirsk team.

## Slide 24: Nuclotron-collider channel diagnostics
The transfer channel will be manufactured by Sigma phi.
Diagnostics will be provided by Instrumentation Technologies.
Shoebox BPM(X/Z) (Libera Spark HL) | 15
Beam loss monitor (Libera BLM) | 10
Fast current transformer (Bergoz 1.2GHz, 5 V/A) + Libera Digit500 | 4
Beam profile monitor (Libera Photon) | 10
* Pictures by P.Leban

## Slide 25: Booster diagnostics layout
D
F
D
F
D
F
D
F
D
F
D
M
M
M
M
M
M
M
M
M
M
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpm
bpm
bpm
bpm
bpm
Vacuum station
F quadrupole
D quadrupole
Magnet
Multipole corrector
BPM
scraper
One quadrant:
Lenses block:
BPM in lenses modules (X/Z) | 24
BPM in magnet modules (X/Z) | 20
Parametric current transformer (Bergoz) | 1
Fast current transformer (Bergoz) | 1
Betatron tune measurements (Q-meter) | 1
Ionization profile monitor | 1
Beam loss monitor | 24
F
Total number of diagnostic devices:
PU(Z)
PU(X)
D
F

## Slide 26: Collider diagnostics
Pick-up near F lenses (X) | 24
Pick-up near D lenses (Z) | 24
Pick-up in straight sections (X/Z) | 20
Parametric current transformer (Bergoz) | 1
Fast current transformer (Bergoz) | 2
Betatron tune measurements (Q-meter) | 1
Ionization profile monitor | 1
Beam loss monitors (for both rings) | 68
BPMs have single plane in arcs and both planes in straight sections.
Current transformers will be installed in warm part of a straight section.
Total number of diagnostic devices in one ring (2 arcs + 2 straight sections):

## Slide 27: Beam Instrumentation in “SPD region”
Stripline BPM, ~1m long, 2 pcs.
Location – next to the Final Focus Quadrupoles. Function – Measurement of the transversal position (X/Y) for opposite directed beams. An additional BPM at IP – indication of the collisions, but small sensitivity.
Electric Kicker, ~1m long, 2 devices (X/Y).
Function – excite beam oscillation in both rings. Betatron tunes Qx,y are measured in each ring separately.
Ionization Profile Monitor+MCP. 3 Detectors for each plane (X/Y).
Function- Measurement of the transversal beam distribution (separately for each ring), definition of beam parameters (emittance).

## Slide 28: Beam Instrumentation in “SPD region”
Equipment composition and layout:
QFF1 – Quadruple magnet, 850mm
Transition «warm-cold» - 500mm
VS – Vacuum shutter, 100mm
BPM – Strip line Beam Position Monitor (X/Y), 1000mm
K – electrical Kicker (X/Y), 1000mm
IPM – Ionization Profile Monitor, 650mm

## Slide 29: Conclusions
Overview of the basic diagnostic systems in Booster, Nuclotron, collider rings and transfer channels, their quantity and placement has been presented.
Thank you for your attention.