## Slide 1
Spectrometer Complex of the IBR-2 High Flux Pulsed Reactor: Current State, Plan of Developments and Recent Results
D.P. Kozlenko
Frank Laboratory of Neutron Physics, Joint Institute for Nuclear Research, 141980 Dubna, Russia

## Slide 2
IBR-2M Spectrometers Complex
Diffractometers:
HRFD, DN-6, RTD, DN-12, FSD, SKAT/Epsilon, FSS
Small Angle Scattering Spectrometer: YuMO
Reflectometers:
REMUR, REFLEX, GRAINS
Inelastic Neutron Scattering Spectrometers:
NERA, DIN-2PI
New Instruments:
DN-6, GRAINS, NRT, FSS
Reconstruction:
REFLEX – SESANS
2011:
11 instruments in operation
2018:
15 instruments in operation
Radiography and Tomography: NRT

## Slide 3
The priority directions of fundamental research :
Nanoscale physics
Physics and Chemistry of Functional Materials
Physics and Chemistry of
Complex Liquids and Polymers
Physics of Soft Condensed Matter

## Slide 4
The priority directions of applied research:
Structural characterization of functional
materials used in different (nano)technologies
Non-destructive control of residual
stresses and internal organization of bulk
materials and products
Texture analysis of geomaterials
and constructional materials
Stainless steel
Zr based alloy

## Slide 5
Diffraction (fundamental research)
HRFD
RTD
Real Time Diffractometer (RTD)
DN-12
Precise structural studies of crystalline materials using reverse TOF method, d/d ~ 0.001
Real time studies
Studies of structure and dynamics of condensed matter under extreme conditions (P ~ 7 GPa)
DN-6
In operation since 2012
Structural studies at ultrahigh pressures (up to 30-50 GPa)

## Slide 6
High Resolution Fourier Diffractometer
t0f = 15 s
I(ω, t) ~ ∫ σ(τ)·R[(τ - t), ω] dτ
0
t
Neutron flux at sample position: 107 n/cm2/s
Resolution at d = 2 Å, 2 =1520: the best achievable resolution d/d = 0.0008
Comparable with advanced instruments in leading neutron centers (SNS, JPARC, ISIS, ILL)
D-spacing range in high resolution mode: 0.7- 4 Å

## Slide 7
Global modernization of the HRFD diffractometer
New Fourier chopper in work position.
Comparison of the old (V-14-10) и new (V-16-10) incident neutron spectra measured using vanadium sample (a). The gain factor of neutron intensity as function of neutron wavelength measured using the standard Al2O3 sample (b).
Neutron guide with vertical parabolic focusing. New mirrors were mounted in old vacuum casing.

## Slide 8
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
3.7E-4 | 5.4E-4 | 6.1E-4 | 6.7E-4 | 7.3E-4 | 7.8E-4 | 8.3E-4 | 8.8E-4
Geometrical contribution to the resolution function for sample of 5x5x5 mm3
Ω = 1.5 Sr
S = 13.5 m2
≈190 elements
Rg: from 0.0004
to 0.0009
Development of New Wide Aperture Back Scattering Detector ZnS(Ag)/6LiF, project within the theme 1122 (NEO KS)

1.813

0.221
8
0.202
7
0.252
6
0.237
5
0.225
4
0.216
3
0.208
2
0.202
1
0.050
0
W, str
Ring
Expected increase in the neutron counting rate is ~ 10 times

## Slide 9
RTD diffractometer for real-time studies
IBR-2:
active core
moderator
Background
chopper
Biological
shielding
6A , 6B beam
splitter
Neutron
guide tube
2D PSD
on rotating arm
Back-scattering
multi-ring detector
SANS
multi-ring detector
90° multi-element
3He detector
Central table,
sample position
Shutter
Neutron flux at sample position: 5106 n/cm2/s
Resolution 2 =156-1700: d/d  0.01 2 =100: d/d  0.05
D-spacing range: 1 - 80 Å
Typical measurement times: 0.5 – 10 min
Parameters comparable with dedicated instruments in other leading neutron centers (SNS, J-PARC, ILL)
T - range:
10 - 300 K
300 – 1000 K

## Slide 10
RTD Sample environment
Refrigerator on goniometer with single crystal inside
Electrochemical cell with vanadium windows
Humidity cell
Thermostat (-40 ≤ T ≤ 100°C)
Further upgrade and development of sample environment is planned
Notes: Продолжались работы по реализация проекта по созданию нового дифрактометра на 6а канале для нейтронографических исследований переходных процессов в реальном времени. Подготовлены рабочие чертежи на изготовление фоновой защиты детекторной системы. Изготовлена и установлена диафрагма нейтронного пучка с регулируемым окном, с линейным перемещением по горизонтальной и вертикальной оси и креплением на выходном фланце нейтроновода RTD.

## Slide 11
DN-12 Spectrometer for Studies of Microsamples
Neutron flux at sample position: 2106 n/cm2/s
Resolution at d = 2 Å, 2 =900: d/d = 0.02
D-spacing range: 0.8 - 13 Å
Pressure range: 0 – 7 GPa
Temperature range: 10 – 300 K
Pressure cell with sapphire anvils
Simultaneous study of crystal and magnetic structure of materials

## Slide 12
Two detectors rings
of DN-12
7
6
1 - first cryostat with coils HTSC magnet (YBCO – superconductor tape); 2 - magnet second cryostat with pressure cell; 4 – current lead; 5 – high pressure chamber; 6 и 7 – detectors DN12; 8 –cryorefrigerator SRDK408S, 9 – cryorefrigerator SRDK101D
n
n
2
1
8
9
4
5
PTH sample environment system (NEO KS FLNP)
3D model of horizontal dry cryostats with split-coil HTSC magnet and cold anvils-insert
Further Developments
Large volume press for high pressure (0 – 30 GPa) and temperature (300 – 2000 K) experiments

## Slide 13
Diffractometer DN-6 for studies of microsamples under extreme conditions
DAC
Neutron diffraction patterns of Cr2O3 sample measured in DAC up to 35 GPa
Second circular detector of DN-6, consisting of 96 separate 3Не counters
Sample volume: 0.01-0.05 mm3
DN-6 diffractometer (operational from 2012
Intensity gain: 12 times (compared to DN-12)
Pressure range: 30-50 GPa
Temperature range: 4– 300 K
Neutron diffraction patterns of Cr2O3 sample measured in sapphire anvil cell up to 4.6 GPa

## Slide 14
The averaged incident neutron flux gain is ~ 6 times.
Multiplication factor compared to DN-12: 35 times
The replacement of the remaining tail part of the mirror neutron guide (17 m) from m = 1 to m = 2 is planned. The expected additional intensity gain is ~ 2 times
Further Development of DN-6: Installation of Neutron Parabolic Focusing Device (7 m length)
Neutron flux at sample position: 3107 n/cm2/s (with focusing device)
Resolution at d = 2 Å, 2 =900: d/d = 0.025
D-spacing range: 0.8 - 12 Å
Pressure range: 0 – 50 GPa
Temperature range: 4 – 300 K
Top level among relevant class of instruments, achievable pressure range comparable with SNAP (SNS)

## Slide 15
FSD
SKAT
EPSILON-MDS
Determination of residual stresses in construction materials, industrial materials and factory-made goods using RTOF method
d/d ~ 0.002 (2 = 90)
Texture analysis of geological materials (rocks, minerals) and constructional materials
Residual stress analysis in geological materials (rocks, minerals)
d/d ~ 0.003-0.004
High resolution due to long flight path L ~ 105 m
Major modernization focused on replacement of neutron guides was finished by 2012
Diffraction (applied research)

## Slide 16
Beam profiles:
analytical calculations and Monte Carlo simulation
Gauge volume of 1.8 mm
FSD: Radial collimators installed in 2016
Loading machine LM-20
Development of the 90-deg scintillation detector ASTRA
Development of new Fourier Chopper

## Slide 17
Installation of the Fourier Stress Spectrometer (FSS) Components at 13 beamline of IBR-2 and first test experiments
The neutron guide, Fourier chopper and detector system of the FSS spectrometer (left). The high resolution spectrum of the iron sample and resolution function (right).
The replacement of the mirror neutron guide with a new one (m = 2, linear convergence), is planned.

## Slide 18
Triaxial pressure device
Parameters:
Paxial < 150 MPa
Pconfining < 70 MPa
Ppore < 70 MPa
EPSILON new opportunities: In situ triaxial pressure experiments including pore pressure

## Slide 19
Small Angle Neutron Scattering and Reflectometry
YuMO - SANS
A study of structural characteristics of nanostructured materials, biological objects, polymers
REMUR
A study of magnetization profile, magnetic and structural properties of layered nanostructures
Reflectometer with polarized neutrons
REFLEX
Reflectometer with polarized neutrons
A study of structural properties of thin films and layered nanostructures
Reconstruction into Spin Echo Small Angle Neutron Scattering Spectrometer
GRAINS
Multifunctional reflectometer
A study of structural properties of liquid and soft matter interfaces

## Slide 20
GRAINS time-of-flight reflectometer with horizontal geometry
Initial beam size | 101 см2
Neutron wavelength | 0.5 – 10 Å
Incident angle | 3 – 25 mrad
q-range | 0.005 – 0.5 Å-1
Angle resolution | 3 - 10 %
Maximal sample size | 5080 mm2
Total neutron flux at sample | 1106 cm-2 s-1
Deflector | Supermirror, m = 2, L = 1 m
Detector | 2D PSD, 3He, 2020 cm2, resolution 2  2 mm2

## Slide 21
Electrochemical cells for studies of processes at the interface metal electrode/liquid electrolyte by neutron reflectometry
Experimental NR curves and fitting results
Project: A system for neutron operando monitoring and diagnostics of materials and interfaces for electrochemical energy storage devices at the IBR-2 reactor (2018-2020)
Polarization analysis development
Polarizer unit upgrade
Installation of analyzer
Working electrode (WE)
Counter electrode (CE)
Reference electrode (RE)

## Slide 22
YuMO Small Angle Scattering Spectrometer
1 – two reflectors;
2 – zone of reactor with moderator;
3 – chopper;
4 – first collimator;
5 – vacuum tube;
6 – second collimator;
7 – thermostate;
8 – samples table;
9 – goniometer;
10-11 – Vn-standard;
12 – ring-wire detector;
13 – position-sensitve edetector "Volga";
14 – direct beam detector.
L1 =17.8 m, L2 =13.2 m
Neutron flux at sample position: 4107 n/cm2/s
Q-range: 710-3 - 0.5 Å-1
Q-resolution: 5-20 %
Typical measurement time: 10 m
Unique dynamic Q-range Qmax/Qmin ~ 90
Fast measurements
Development of new incident beam collimator, modernization of detector system, sample environment

## Slide 23
REMUR - Reflectometer with polarized neutrons
1 - Moderator
2 - Double Disk Background Chopper
3 - Collimators
4 - Cross-type Collimator
5 - Small-Angle Scattering Mode Polarizer
6 - Adjustable Platforms
7 - Spin-Flipper
8 - Variable Diaphragms
9 - Sample Position
10 - Electromagnet
11 - Fan Polarization Analyzer
12 - Position-Sensitive Detector
13 - Control and operative visualization/analysis
14 - Data Acquisition
15 - Data Transfer
Neutron flux at sample position: 3104 n/cm2/s
-range: 0.9 - 10 Å-1
2 range: 1-100 mrad
/: 2 %
Reflection coefficients from
12[Fe(35Å)Cr(4.4 Å)/Gd(50 Å)]

## Slide 24
Modernization of REMUR: installation of new polarization analyzer
Isotope-identifying neutron reflectometry (IINR) project, 2015-2018
HPGe detector for gamma-channel

## Slide 25: Reconstruction of the REFLEX Reflectometer into Spin Echo Small Angle Neutron Scattering Spectrometer
Time gradient magnetic fields
End part of a new neutron guide assembled at the 9th beamline.
Incident neutron wavelength distribution spectra on 9th beamline at the previous instrument configuration (black) and with a new neutron guide (red).
Intensity: I  106 n/s/cm2
SE length interval: ZTGF = 100 Å  5000 Å

## Slide 26
Experimental test of SESANS setup at 9-th beamline of the IBR-2 reactor
X
Y
Z
Spin rotator
Fig. 1. A) Wavelength dependence of the polarization Pz of the outgoing neutron beam for the symmetric setup without sample.
B) NSE signals observed for indicated wavelength bands (1,2,3,4,5).
1
2
3
4
5
1
2
3
4
5
A
B

## Slide 27
Inelastic Neutron Scattering
DIN-2PI
NERA
Vibrational spectra of molecular crystals
Lattice dynamics, atomic dynamics of liquids and amorphous materials
DN-12
Vibrational spectra of hydrogen-containing materials under pressure up to 5-10 GPa

## Slide 28
NERA – Inelastic Neutron Scattering Spectrometer in Inverted Geometry
1 - Moderator
2 - Background Choppers
3 - Ni Guide Tube
4 - Detector for High Intensity Diffraction
5 - Detector for High Resolution Diffraction
6 - He3 Detectors (INS and QNS)
7 - Single Crystal QNS Analyzer
8 - Pyrolytic Graphite INS Analyzer
9 - Be-Filters
10 - Collimators
11 - VME control and operative visualization/analysis
12 - VME Station (OS/9)Data Acquisition
13 - EtherNet Data Transfer
Neutron flux at sample position: 4106 n/cm2/s
Energy transfer range: 0-500 meV
NERA integral flux gain: 1.984
Major modernization focused on replacement of neutron guide was finished by 2012. The fabrication of parabolic focusing device is in progress.

## Slide 29
DIN-2PI – Inelastic Neutron Scattering Spectrometer in Direct Geometry
Neutron flux at sample for E0 = 10 meV: 2103 n/cm2/s
2 range: 5-1350
Incident neutron energy range: 1-30 meV
Energy resolution: E/E = 4-10 %
Installation of neutron concentrator gives an order of magnitude increase of neutron flux at sample position
2006
2011
FWHM of incoherent QENS peak in liquid Ga

## Slide 30
 | Planned | Achieved
Neutron flux at the sample position | 106 n/cm2/s | 5106 n/cm2/s
Incoming beam apertures | 50, 20, 10, 5 mm | 50, 20, 10, 5 mm
L/D ratio for given apertures | 194, 485, 970, 1940 | 194, 485, 970, 1940
Fields of view for given L/D | 289, 230, 210, 200 mm | 200 mm
Spatial resolution | 100- 200 m (scintillator 100 -200 m thickness) | 150-250 m
CCD-camera parameters | Active pixels: 4008×2672 Pixel size: 9×9 m Image area 36×24 mm Digitization: 12 bit Cooling (Peltier) to -25 C | 
Parameters
Spectrometer for Neutron Radiography and Tomography

## Slide 31
Development of Neutron Radiography and Tomography Spectrometer: Installation of Precision Two Mirror Optical System

## Slide 32
 | Number of mirror channels | Length | Radius of curvature
mod1 | 20 | 1 m | 7.16 m
mod2 | 20 | 2 m | 14.32 m
mod3 | 30 | 1 m | 7.16 m
mod4 | 30 | 2 m | 14.32 m
Bender modifications (m = 2)
M.V.Avdeev, R.A.Eremin, V.I.Bodnarchuk, I.V.Gapon, V.I.Petrenko, R. Erhan, A.V. Churakov, D.P.Kozlenko,
J. Surf. Investigation. (2018) in press.
Spectrum calculations
Flux distribution
32 m
Beamline 10A
T = 100 K
A concept of new small angle scattering and imaging spectrometer at 10-A channel of IBR-2
ССD camera based detector
Beam size | 5050 мм2
Neutron wavelength range: | 0.5 – 10 Å
q-range: | 0.003 – 0.5 Å-1
Angular resolution | 5 - 20 %
Maximum sample size | 205050 мм2
Incident neutron flux at the sample | 2.0106 см-2 c-1
Detector system | PSD, 3He, 6464 см2, resolution 55 мм2
Expected parameters

## Slide 33
Temperature of moderator | 30 K | 100 K | 300 K
Before bender | 1.0e9 | 4.3e8 | 1.8e8
After bender | 3.9e8 | 8.5e7 | 1.4e7
Sample position (collimation length 1 м) | 2.3e8 | 5.6e7 | 1.0e7
Sample position (collimation length 10 м) | 7.4e6 | 2.7e6 | 7.2e5
Total flux calculations
(flux density on moderator 1012 cm-2 s-1)
Total flux measurements
(monitor PSD)
Temperature of moderator | 30 K | 300 K
Before bender | > 1.0e8 | ~5.0e7
After bender | ~5.0e7 | ~8.0e6
30 K – working mode (flux at sample > 106 cm-2 s-1)
300 K – mode for high-scattering systems
(flux at sample > 105 cm-2 s-1)
Bender Tests at 10A Beamline
Mirrotron (Budapest, Hungary)
L = 2 m, R = 14 .3 m, N = 20 ( m = 2)
T = 300 K
Before bender
(detector overload)
After bender
TOF
T = 30 K
Before bender
(strong detector overload)
After bender
TOF

## Slide 34
Thank You for Your Attention!