## Slide 1: Radiation hardness tests of GaAs and Si sensors at JINR S.M. Shakour JINR, Dubna, Russia 2017

## Slide 2: Gallium arsenide as a detector material
Material | main charge carriers | electron drift length | bulk resistivity | active sensor thickness | intrinsic noise
LEC SI-GaAs | holes | 0.3-0.5 mm | <2*108 Ω*cm | <300 μm | high
GaAs:Cr | electrons | 0.7 – 2 mm | ~109 Ω*cm | up to 1 mm | low
20%
GaAs is a well-known semiconductor, second widespread after silicon
Limited use in particle detection because of low resistivity, low CCE and high intrinsic noise
New modification of GaAs, compensated by Cr (GaAs:Cr), has been invented in Tomsk State University in 2000-2005
  suitable for detector construction
  radiation hard
  Z(GaAs)~ 32 vs Z(Si)=14 →
higher photon detection efficiency

## Slide 3: Two types of GaAs:Cr detectors
'Resistive' GaAs:Cr
  resistivity ~109 Om*cm
  active thickness up to 1 mm
  electron drift length up to 2 mm
πν junction structure
  active thickness is determined by πν junction (~0.1-0.2 mm depending on Ubias)
  resistivity and CCE is Ok
ν-type GaAs:Cr
π-type GaAs:Cr
ν-type GaAs:Cr
n-type GaAs

## Slide 4: Summary of previous measurements
GaAs:Cr detectors based on πν junction structure may have the same or better radiation hardness comparing with the 'resistive' GaAs:Cr
πν junction structures were never tested systematically in electron beams.
We never reported (= never systematically studied) the dependence of the radiation degradation on the irradiation rate

## Slide 5: GaAs:Cr and Si sensors
4
2
3
1
N | Type | Holder | Size, x y z, mm | Sensitive area, mm2
1 | GaAs:Cr barrier, n+-π-ν-n(radiation hard) | Plastic | 5x5x0.3 | 5x5
2 | GaAs:Cr high resistive | PCB | 5x5x0.3 | 4.5x4.5
3 | Si n-type | Plastic | 5x5x0.3 | 4x4
4 | Si n-type rad-hard sensor from USCS | PCB | 10x10x0.4 | 4x4
1

## Slide 6: LINAC-800
800 MeV electron linear accelerator (LINAC-800) is under construction at JINR. 20 MeV beam channel was used for sensors irradiation. Beam parameters:
bunch current up to 10 mA,
duration 2 ms,
frequency from 1 to 10 Hz.
Set up our experiments

## Slide 7: Irradiation setup
Keithley 6517
Beam pipe
Collimator
Sensors
Faraday cup

## Slide 8: Irradiation control
The radiation sensitive film was placed behind the sensor
to control the absorbed dose, and the uniformity of electron flounce during irradiation.
48 kGy
194 kGy
Sensor
5mm
5mm
5mm
5mm

## Slide 9: GEANT 4 simulation
Sensor | Holder | Thickness, um | Energy deposit by one e, MeV | Absorbed dose to charge in Faraday cap, kGy/uC | Error
GaAs | Plastic | 300 | 0.2218 | 5.56 | 0.04
GaAs | PCB | 300 | 0.2238 | 5.61 | 0.04
Si | Plastic | 300 | 0.1064 | 6.09 | 0.06
Si | PCB | 300 | 0.1066 | 6.10 | 0.06
The electron transport through the irradiation setup was simulated by GEANT4 in order to obtain the ratio of registered by Faraday cup charge to absorbed dose in the sensor.

## Slide 10: CCE(Charge collection efficiency ) measurement set up
electrons from Sr90 source well collimated and triggered by 2 scintillators. It allows to cut and measure signal only from electrons passed throw the sensor with energy from 1 to 2.2 MeV which is close to MIP electrons

## Slide 11: Installation for CCE and I-V(Volt - Ampere) measurement
All the sensors measured at room temperature, except some measurements at low temperature for Si

## Slide 12: Results I-V Si
I-V characteristics of the normal n-type Si (N3) and radiation hard USCS Si (6888).
The dark current increased almost 4 orders of magnitude after the absorbed dose of 1.5 MGy

## Slide 13: Results I-V: πν junction structure and HR GaAs:Cr
I-V characteristics of the sensors: high resistive GaAs:Cr (left), πν junction structure
GaAs:Cr(right)
Resistivity of GaAs:Cr sensors is about 2-5*109 Ohm*cm, the dark current increased only 3 times for the high resistive and 4-5 times for the barrier GaAs:Cr.

## Slide 14: MIP of non irradiated GaAs:Cr sensors
MIP spectrum from Sr90 of the HR GaAs (left) and πν junction structure GaAs sensors (right)

## Slide 15: GaAs MIP signal vs Dose
0kGy
48kGy
190kGy
350kGy
550kGy
950kGy
1150kGy
1350kGy
1550kGy
Mean : 34.93
Mean : 22.21
Mean : 14.05
Mean : 13
Mean : 8.5
Mean : 4.64
Mean : 4.99
Mean : 4.97
Mean : 4.95

## Slide 16
GaAs pedestal vs Dose
1550kGy
48kGy
0kGy
In GaAs:Cr the CCE drops 5 times even at a dose 0.5 MGy, but then it decreases slowly,
keeping about 10% of the initial CCE after dose 1.5 MGy. The width of the pedestal in a
GaAs remains practically unchanged.

## Slide 17
UCSC Si 6886 MIP vs Dose
493kGy Sig
0kGy
493kGy Sig+Ped
493kGy Ped
The pedestal expands and the signal from pedestal separation becomes difficult after doses 0.5MGy

## Slide 18
N-type Si N5 MIP vs Dose
0kGy Sig
104kGy Sig
208kGy S+P
274kGy S+P
383kGy S+P
492kGy S+P
602kGy S+P
T =22°C
T =20°C

## Slide 19
Si vs GaAs 550 kGy
Si N5; 492kGy ; T= 16°C
GaAs N4 barrier; 550kGy ; T =22°C
PED
SIG
A slight cooling of the sensor Si for 5°C significantly improves the MIP-signal and pedestal separation, but it is still worse than GaAs:Cr.

## Slide 20
Si vs GaAs 1550 kGy
Si 1260 kGy ; T =22°C
GaAs 1550 kGy ; T =22°C
PED
SIG
After 1.5 MGy the MIP-signal and the pedestal still separated for GaAs:Cr

## Slide 21
Si cooling
Si 602 kGy ; T =20°C
Si 602 kGy ; T =-1°C
PED
SIG

## Slide 22
Si cooling
Si 1260 kGy ; T =22°C
Si 1698 kGy ; T =-0.7°C
PED
SIG

## Slide 23
GaAs CCE vs Dose
(GaAs)
(GaAs)

## Slide 24
Si CCE vs Dose
Due to increasing of the dark current the pedestal considerably broadened and the measurement becomes difficult for doses higher 0.5 MGy.

## Slide 25
GaAs:Cr and Si sensor resolution.
2σ-criterion was applied for a correct comparison of Si and GaAs sensor performance.
K is the ratio of events departed greater than the total number 2σ from pedestal to the total number of events in MIP-spectrum.
The greater K, the better the signal and pedestal are separated in the spectrum.
GaAs barrier
Si N5

## Slide 26
GaAs CCE-Vbias dependence.
Chart: E, kV/cm
| 0.39999999999999997 | 1 | 2 | 4 | 5.9999999999999991 | 8 | 10 | 11.999999999999998 | 13.999999999999998 | 16
9639#25N1 | 0.41188 | 0.45438 | 0.48425 | 0.52554 | 0.55978 | 0.57872 | 0.60787 | 0.62584 | 0.64672 | 0.66858
N7 | 0.40976 | 0.45502 | 0.48637 | 0.50779 | 0.52236 | 0.53649 | 0.55481 | 0.57071 | 0.58594 | 0.6014
G5557_1.2N4 550 kGy | 0.033196 | 0.056878 | 0.11206 | 0.15837 | 0.18374 | 0.20552
G5557_1.2N6 1550 kGy | 0.013109 | 0.030236 | 0.050957 | 0.069776 | 0.081405 | 0.093034
N23 1550 kGy | 0.024104 | 0.047237 | 0.059289 | 0.065898 | 0.076589
For the irradiated GaAs sensors CCE increases significant with the field strength up to 20 kV / cm,
while in the non-irradiated sensors saturation begins at 1 kV/cm
GaAs HR
GaAs barrier
GaAs barrier
GaAs barrier
GaAs HR

## Slide 27
CCE measurements for GaAs and Si under neutron irradiation from IBR 2

## Slide 28: Summary
Irradiation of the high-resistive and barrier GaAs:Cr sensors and the normal and radiation-resistant n-type Si sensors by 20 MeV electron beam of was performed.
Radiation damage in GaAs and Si are different: increasing of the dark current in Si and dropping signal in GaAs.
A significant difference between two types of GaAs:Cr sensors and two types of Si is not found.
2-σ criterion was applied to separate the signal from the pedestal.
After dose 1.5 Mgy in GaAs:Cr signal magnitude remains about 10% of initial, when Si is not possible to use at room temperature.