## Slide 1: Advances in the development of a Vacuum Insulated Tandem Accelerator and its applications
Budker Institute of Nuclear Physics, Novosibirsk, Russia Novosibirsk State University, Novosibirsk, Russia
Sergey TASKAEV, Alexander IVANOV,
BINP BNCT team: Evgenii BERENDEEV, Marina BIKCHURINA, Timofei BYKOV, Dmitrii KASATOV, Anna KASATOVA, Iaroslav KOLESNIKOV, Alexey KOSHKAREV, Alexandr MAKAROV, Georgii OSTREINOV, Alexey PONEDELCHENKO, Pavel PONOMAREV, Sergey SAVINOV, Ivan SHCHUDLO, Evgeniia SOKOLOVA, Igor SOROKIN, Tatiana SYCHEVA, Iuliia TASKAEVA, Gleb VERHOVOD

## Slide 2
Achievements – significant
Applications – numerous
Facility is a state-of-the-art device comprised
the Vacuum Insulation Tandem Accelerator (VITA)
a solid lithium target
a Beam Shaping Assembly

## Slide 3
The original design tandem accelerator, which was named as Vacuum-Insulated Tandem Accelerator (VITA), has a specific design that does not involve accelerating tubes, unlike conventional tandem accelerators. Instead of those, the nested intermediate electrodes (5) fixed at single feedthrough insulator (8) is used, as shown in Figure.
The advantage of such an arrangement is moving ceramic parts of the feedthrough insulator far enough from the ion beam, thus increasing the high-voltage strength of the accelerating gaps given high ion beam current.
A consequence of this design was also a fast rate of ion acceleration—up to 25 keV/cm.
VITA: 1 – H– source, 2 – magnetic lenses,
3 – accelerator, 4 – cryogenic pump, 5 – intermediate electrodes, 6 – high-voltage electrode, 7 –gas stripper, 8 – insulator, 9 – high-voltage power supply, 10 – turbomolecular pumps.

## Slide 4
For the compactness of the accelerator, the average electric field strength on the insulator was chosen to be 14 kV/cm, which is 1.5 times higher than the recommended one.
This led to breakdowns along the surface of smooth ceramic insulators with a height of 73 mm, from which the feedthrough insulator was assembled.
Such breakdowns occurred approximately once every 10–20 min; they did not lead to a decrease in the electric strength of the accelerator, but required 15 s to restore the parameters of the ion beam.
To eliminate breakdowns along the vacuum surface of insulators, smooth ceramic insulators were replaced with corrugated insulators.
Advance # 1: elimination of high-voltage breakdowns

## Slide 5
Advance # 2: wire scanner OWS-30 (D-Pace, Canada) is used to control the position and size of the ion beam at the entrance to the accelerator
The phase portrait of the ion beam
The dependences (a) of the current I, (b) the cross sectional area S, (c) the current density j of the ion beam on the residual gas pressure P
source of negative hydrogen ions,
cone diaphragm,
vacuum lamp,
turbomolecular pumps,
magnetic lenses,
movable diaphragm,
OWS-30 wire scanner,
cooled diaphragm,
first electrode of the accelerator,
vacuum tank of the accelerator,
metal rings,
leak valve.

## Slide 6
Advance # 3: The position and size of the ion beam in the accelerator are controlled by two pairs of video cameras overseeing the input and output diaphragms of the external accelerating electrode
The presence of a gas stripper in a tandem accelerator is often considered a disadvantage. We managed to neutralize the disadvantages and turn the use of a gas stripper into an advantage. The additional gas flow makes it possible to visualize the beam for diagnostics of its position and size and improves the high-voltage strength of the accelerating gaps.

## Slide 7
Advance # 4: The position and size of the ion beam inside the gas stripper is monitored with Celestron Ultima 80-45 telescope looking axially through a cooled mirror.

## Slide 8
Advance # 5: The magnitude of the argon ion beam current was measured by mass spectroscopy using a bending magnet and a cooled diaphragm; it is 2000 times smaller than the proton beam current.
The reliability of the measurement is provided by visualization of an argon ion beam on the surface of a lithium target, as confirmed by an experiment with increased gas injection and an estimate of the possible contribution of the proton beam.
Such a small value of the current of the argon ion beam poses no danger either as an additional heating of the lithium target or as an additional load of a high-voltage power supply and therefore does not require the previously proposed suppression means.

## Slide 9
Advance # 6: No appreciable influence of the spatial charge on the proton beam transport.
I, mA
r, mm

## Slide 10
Advance # 7: 9 methods for measuring the proton beam profile implemented. Phase portrait of a beam of protons and neutrals measured.
The proton beam has a characteristic size of 11 mm,
a divergence of 1.5 mrad, and a normalized emittance of 0.23 mm mrad.

## Slide 11
High power DC proton/deuteron beam (20 kW):
Monochromaticity and stability: 0.1%
Energy: ranges from 0.6 MeV to 2.3 MeV
Current: ranges from 1 pA to 10 mA (10 orders!)
High flux neutron beam (2 1012 s-1):
thermal (D2O or Plexiglas moderator)
epithermal (MgF2 moderator)
exclusively epithermal (no fast and thermal)
over-epithermal
monoenergetic (kinematic collimation)
fast
Bright source of photons:
478 keV - 7Li(p,p’)7Li
511 keV – 19F(p,e+e-)16O
-particles in 7Li(p,) and 11B(p,)
positrons in 19F(p,e+e-)16O
RESULT:

## Slide 12
Application # 1: Boron Neutron Capture Therapy
Neutron source VITA meets the requirements of BNCT to the greatest extent
The neutron source VITA commercialized by TAE (CA, USA): 1st source is installed in new BNCT Center at Xiamen Humanity Hospital
in Xiamen, P.R. China in 2020. It is planned to start treatment in early 2022;
2nd source will be made for National Oncological Hadron Therapy Center (CNAO) in Pavia, Italy;
3rd source will be made for National Medical Research Center of Oncology in Moscow, Russia;
- I hope the list will go on
Visit of the Prime Minister of the Russian Federation M. Mishustin to BINP March 05, 2021

## Slide 13
Application # 1: Boron Neutron Capture Therapy
Our new results:
A small-sized detector for measuring boron dose and gamma-ray dose is developed and used [JINST 16 (2021) P01024];
New method for measuring the sum of the fast neutron dose and the nitrogen dose is proposed and implemented [Radiation Research 196 (2021) 192-196];
New method for in situ measuring the lithium layer thickness is proposed and implemented [JINST 15 (2020) P10006];
The neutron yield of the 7Li(p,n)7Be reaction in lithium target was measured [Biology 10 (2021) 824];
A number of new boron delivery drugs have been tested [Intern. J. Molecular Sciences 22 (2021) 7326; Intern. J. Radiation Oncology, Biology, Physics; Pharmaceutics 13 (2021) 1490; Chemical Phys. Letters; …]
Our patent on new boron delivery drug, jointly with colleagues from Moscow, [RU2729458, WO 2020/246913] is recognized as the best Russian invention of the 21st century.

## Slide 14
Application # 2: Fundamental knowledge
7Li(p,p‘g)7Li reaction cross-section

## Slide 15
Application # 2: Fundamental knowledge
7Li(p,p‘g)7Li reaction cross-section
[Nucl. Instrum. Methods Phys. Res., Sect. B, vol. 502, pp. 85-94, June 2021]

## Slide 16
Application # 2: Fundamental knowledge
7Li(p,p‘g)7Li reaction cross-section
[Nucl. Instrum. Methods Phys. Res., Sect. B, vol. 502, pp. 85-94, June 2021]

## Slide 17
Application # 2: Fundamental knowledge
7Li(p,p‘g)7Li reaction cross-section
[Nucl. Instrum. Methods Phys. Res., Sect. B, vol. 502, pp. 85-94, June 2021]

## Slide 18
Application # 2: Fundamental knowledge
478 keV photon yield from a thick lithium target

## Slide 19
Application # 2: Fundamental knowledge
478 keV photon yield from a thick lithium target
[Nucl. Instrum. Methods Phys. Res., Sect. B, vol. 502, pp. 85-94, June 2021]

## Slide 20
Application # 2: Fundamental knowledge
478 keV photon yield from a thick lithium target
[Nucl. Instrum. Methods Phys. Res., Sect. B, vol. 502, pp. 85-94, June 2021]

## Slide 21
Application # 2: Fundamental knowledge
478 keV photon yield from a thick lithium target
[Nucl. Instrum. Methods Phys. Res., Sect. B, vol. 502, pp. 85-94, June 2021]

## Slide 22
Application # 3: ITER
[Fusion Eng. Des., vol. 168, 112426, March 2021; Nuclear Fusion (2021)]
Boron carbide ceramics and steel 316L(N)-ITER Grade for ITER (Cadarache, France) were tested by thermal and fast neutrons

## Slide 23
Application # 4: CERN
The neutron source is planned to be used for radiation tests of fibers of the laser calorimeter calibration system of the Compact Muon Solenoid electromagnetic detector developed for the High-Luminosity Large Hadron Collider in CERN

## Slide 24
Application # 5: Proton microscope
elemental analysis of the surface of materials by backscattered protons
Signals of the alpha-spectrometer upon irradiation with 1 MeV protons of the freshly deposited lithium layer of the target (1) and the lithium layer after deliberate admission of atmospheric air into the target unit for 10 min (2) and for 1 h (3)
We determine that on the surface of the freshly deposited lithium layer there are lithium oxide of 2.4 ×1016 atoms/cm2 density
(~ 6 nm of lithium oxide of crystal density) and carbon over it of 6×1015 atoms/cm2 density (~ 0.5 nm of carbon of crystal density)
150 nm boron carbonitride film on 0.5 mm silicon substrate

## Slide 25
Application # 6: Positrons source

## Slide 26
Our plans:
to make a unique proton micro- nanoscope;
to test new boron delivery drugs for many researchers groups;
to study the radiation resistance of optical fibers developed for the High-Luminosity Large Hadron Collider in CERN jointly with Saclay Nuclear Research Center (France);
to study the activation of materials developed for ITER (France) jointly with Russian ITER team;
to study the energy and angular characteristics of the 11B(p,) reaction products jointly with BINP detectors laboratory;
to characterize the epithermal neutron field jointly with the Laboratory of Subatomic Physics and Cosmology CNRS-IN2P3, Grenoble-Alpes University (France) and the Laboratory of Micro-Irradiation, Metrology, and Neutron Dosimetry, IRSN (Cadarache, France);
to develop boron imaging technique jointly with the University of Pavia (Italy), and the Due2lab s.r.l. (Parma, Italy);
to produce ultracold neutrons and focus them jointly with National Laboratory of Frascati (Italy);
to realize new ideas in VITA, lithium target, and BSA for an even better neutron source for BNCT clinics;
…

## Slide 27
CONCLUSION
Neutron source based on VITA and lithium target is in operation now
This source produces:
DC proton/deuteron beam (0.6 MeV – 2.3 MeV; 1 pA – 10 mA)
Neutrons – thermal, epithermal, over-epithermal, monoenergetic, fast
Photons – 478 keV, 511 keV
-particles – 7Li(p,), 11B(p,)
Positrons – 19F(p, e-e+)16O
We are open for joint research
This work was supported by the Russian Science Foundation (grant number 19-72-30005)

## Slide 28
Thank you for
your attention!