## Slide 1: Temperature and Humidity control inside STS box
Marcel Bajdel
Goethe-Universität (Frankfurt)

## Slide 2
Sensors for HEP
Are there any commercial humidity sensors suitable for HEP?
Industrial high-precision dew point transmitters (e.g. Michell, Vaisala)
  High performance, good accuracy, down to dew points-120°C
  Instrumentation/readout/pump need to be placed in area with low/no radiation
  Overall system is bulky and expensive (5k €)
  No distributed sensing
Capacitive relative humidity sensors
  Miniaturized
  SHTxx not radiation resistant but as per CERN MK33-W sensors from IST are radiation hard
  Magnetic field (?)
FBG/LPG based sensors
  Relatively small
  Magnetic field insensitive
  Radiation hard
  Difficult to integrate in a hardware interlock
  Highly dependent on temperature changes
  Readout device is expensive (20k €) but it can be used for hundreds of sensors
Forum Tracking Detector Mechanics

## Slide 3
STS humidity and temperature
ASICs internal T sensor,
GBT SCA T sensor
Sniffing tube and readout system
Commercial capacitive sensor (SHT85 and P14)
Fiber Bragg Grating Fiber Optic Sensors

## Slide 4
ASIC internal T sensor
Low precision (4oC)
Can’t be readout during data taking
~14 000 T sensors – averaging per FEB BOX(?)
Easy integration with DCS
ASICs internal T sensor,
GBT SCA T sensor
Cooling block temp
from -30oC to 20oC

## Slide 5
T FOS
Peak reflectivity
Pe – photo-elastic constant, ϵ - strain induced on the fiber, α – thermal-expansion coefficient, ζ – thermo-optic coefficient
Reflectivity of the T FBG
FBG is in a strain-free conditions and strain effect is compensated, then:

## Slide 6
RH FOS
αt - thermal expansion coeffiecient, αrh – moisture expansion coefficient
Reflectivity of the RH FBG
G.M. Berruti, CERN-THESIS-2015-321

## Slide 7
Testing
2x single RH/T sensor with different humidity sensitivity
5xRH/T arrays to be packaged with carbon fiber packaging
Radiation hard, custom packaging
Not radiation hard, to be packaged
Setup
Climatic chamber for temperature control
Saturated salt solution + N2 for humidity control
SHT85 and Michell Chiller Mirror for calibration
Readout (interrogator) integrated in the DCS

## Slide 8
RH and T FOS
RH FOS response in comparison to the Michell Chilled Mirror

## Slide 9
SHT85
TFOS 1-2% difference compared to PT1000 after nonlinear fitting

## Slide 10
SHT85
SHT85 1-2% RH error down to 0oC

## Slide 11
Sniffing system
Air will be sniffed from the STS box and measured ~25 meters away:
High precision of dew point measurement (1-2 oC) down to -100oC
Perfect for implementation in the interlock system
Two such devices will be tested in the Thermal Demonstrator

## Slide 12
Conclusions
 | Radiation hardness | Magnetic field resilience | Miniaturized | Precision | Integration efforts | R&D needed | Price
Fiber Optic Sensors |  |  |  | Probably 1-2% RH(?) ~0,1 oC T | High | Months - years | 20-30 k EUR + 50-500EUR per sensor
SHT85 |  |  |  | 1-2% RH down to 0oC ~0,1oC T | Low |  | 30 EUR per sensor + readout (?)
MK33-W |  |  |  | 1-2% RH down to 0oC | Low | Months | 30 EUR per sensor + readout (?)
Sniffing system |  |  |  | 1-2oC DP down to -100oC | High |  | 5-6k EUR per device
ASIC internal T |  |  |  | Low 4oC |  | Calibration |