## Slide 1
Optical quality assurance procedures for the components of the CBM STS Detector
Evgeny Lavrik,
Eberhard-Karls Universität Tübingen
for CBM Collaboration

## Slide 2: Overview
Optical inspection setup
  Overview
  STS Sensors
  Inspection setup and inspection principles
  Capabilities
  QA software
  QA Database
  Machine learning approach for QA
Inspection of cables
Inspection and metrology of sensors on the ladder
Summary

## Slide 3: Silicon Tracking System (STS) detector
Compact detector built out of ~900 silicon microstrip sensors
8 layers of sensors
4 sensor size types
2 sensor vendors
STS detector without
thermal insulation

## Slide 4: STS Sensors
Double-sided micro strip Si sensor
0° (n-side), 7.5° (p-side) stereo angle
58 µm strip pitch
1024 strips per side
6.2x12.2, 6.2x6.2, 6.2x4.2, 6.2x2.2
cm2 form factors
2 manufacturers (CIS, Hamamatsu)

## Slide 5: Optical inspection setup
Flexible design to support inspection of different objects (different sensor size-types from CIS and Hamamatsu, sensor micro-cable inspection), other metrology and microscopy tasks
Low hardware dependence, adaptable to almost any hardware
Configurable QA procedures as plug-ins
Report building, storage, viewing and editing
Constant improvement in performance (inspection times 1 hour -> 30 min per sensor side, faster with faster camera) and inspection quality

## Slide 6: Inspection setup capabilities
Possible to detect:
  Dust particles and other foreign objects on the surface
  Scratches
  Single element integrity
    bias resistors
    strips
    pads
    guard ring
  Sensor edge defects & parallelity
  Possible any deviation from
  clean pattern (pattern/texture matching)
  Sensor warp inspection
Setup evolved to the metrology station
Edge profile
Recognized surface scratch

## Slide 7: XY- and Z-stage calibration
Calibration movement pattern
Scan movement pattern
Extract Stretching, Conversion, Rotation matrices
Images taken and sent to defect analysis
XYZ Stages has been characterized with a calibration plates
XYZ Positioning error ± 2.5 µm
XYZ Repeatability error
± 1 µm

## Slide 8: Optical axis calibration
Camera properties.
Here exposure
Height measurement parameters
Here depth of field, related to HWHM
Standard procedures to characterize optical system, HW independent

## Slide 9: Optical axis calibration
Software parfocality correction
Pixel-to-µm conversion ratio
Standard procedures to characterize optical system, HW independent

## Slide 10: Height measurements
Contactless, thus non-damaging method to measure object heights
Uses Z- and Focusing Stage of the inspection setup, depending on precision
Most focused value extracted from Lorentzian fit
Source, FFT-transformed image
and magniture of all complex pixels
at different focus values
Source
Changing focus
Changing Z-position

## Slide 11: Height measurements
Height map (warp) of a CBM06C6 sensor
Measured thickness, µm |  |  | 
CIS |  | Hamamatsu | 
bulk | strips | bulk | strips
303 | 311 | 331 | 340
Every 10 µm of silicon generate ~100 e- per MIP,
The expected read-out signals could be predicted
Quality criterion
Measuring most focused values at different positions allows to:
  Sensor warp measurements
  Sensor thickness measurements
  Other metrology procedures

## Slide 12: Sensor Warp Measurements
Calibration against a certified Mitutoyo gauge block set yields precision of ~1µm
Performance optimizations with adaptive method, speedup of 8-20 times
Meas step, mm | Adaptive method, hrs | Straightforward, hrs
1 | 13.75 | 106.12
2 | 3.43 | 26.53
3 | 1.5 | 11.8
4 | 52 min | 6.63
5 | 33 min | 4.2
Warp measurement time
of a 6x6 sensor

## Slide 13: Sensor Warp for 20 sensors
Values of warp observed between 15 and 80 µm across prototype sensors,
seemingly no dependence (yet?)

## Slide 14: Sensor cutting edge
Cutting edge quality estimation
by looking at the deviation from fitted edge
≤ 20µm
Cutting edge parallelity
by looking at the distance from the alignment mark to the fitted edge at all 4 corners

## Slide 15: Inspection reports
Inspection results with detected defects, different annotation
Here a baby sensor for demonstration purposes

## Slide 16: Inspection reports
Single ROI with defects detected, defect information (class, center of mass, area, etc.)
Operators workplace with navigation, editing etc. functionality

## Slide 17: Database
Reports formed during analysis to be stored in Database
Centralised data storage for CBM – FairDB
1 full inspection is 25.8 GB per 6x6 sensor (n and p sides, lossless PNG)
Up to 50 TB of images needs to be stored.
  Tape storage gStore from GSI for archiving
  HPC clusters at GSI for “hot“ storage
Database interfaces were developed for optical QA

## Slide 19: Defect detection in a context
Problem: In direct light all surface defects appear as dark objects
How to differentiate between them?
Example: Dust particles cover strips and interstrip area,
Scratches appear mostly only on strips.
Knowing the context of defect affects its severity weighting
Currently identified by by pattern matching. Not universal.

## Slide 20: Identifying the context
Pattern matching does its job well, but needs to be supervised
Adjusting the tresholds and matching scores during production phase is a bad design
Idea: augment it with machine learning methods
E.q. a classification neural network, which adapts itself to everchanging global data.
The machine vision and machine learning enjoys a lot of academic interest in the last time with new ideas, models and software frameworks being constantly published
A „Darknet“ framework with a model „You only look once“ has been taken into consideration for its performance and relative ease of use

## Slide 21: Applying it to our task
Train on an imageset with features marked
~ 12 Hours on OEM GTX 745
Synthetic test shows a good output
Here shown is the detection of features and electrical components of a Hamamatsu sensor

## Slide 22: Defect Classification

## Slide 23: Microcable inspection motivation: Low Cable Yield at LTU

## Slide 24: Microcable structure
Cables leave the LTU production factory with signal layers embedded between mesher spacers and bundled into cable stacks.
Only place for optical QA – production place

## Slide 25: Microcable inspection principles shown during last CBM collaboration meeting

## Slide 26: Sensor inspection and metrology on the ladder assembly tool
Assembly tool CAD
Tool built in
the optical inspection setup
See talk of Uli Frankenfeld on
assembly procedures after lunch

## Slide 27: Metrology of the sensors on the ladder
CBM06C6
n-side
CBM03C6 2009
p-side
Measurement | Value, µm
Width (top) | 3919
Height (top) | 37
Depth (top) | 536
Depth (bottom) | 297
Sensor overlap measurements
Euler angle | Value, °
φ | 0.011
θ | 0.247
ψ | 0.225
Sensor to sensor angular misalignment
Sensor configuration
Edge | Unparallelity, µm | 
 | Left Sensor | Right Sensor
Top | < 1 | -7.5
Bottom | -2.6 | 3.3
Right | < 1 | N/A
Left | 4.7 | 2
Cutting edge unparallelity

## Slide 28: Sensor spatial orientation

## Slide 29: Sensor warp on the ladder
Max warp 24 µm
Max warp 19 µm
Check if additional mechanical stress is created after glue hardening...
Seems not to be the case.

## Slide 30: Summary
Flexible optical inspection and metrology setup to inspect different components of STS and other tooling
Low hardware dependence, standard calibration routines
Advanced and optimized analysis methods for:
  Sensors (CiS and Hamamatsu)
  Microcables (LTU, KIT to come)
  Ladder assmbly tool (GSI, JINR to come)
  ... anything