## Slide 1: Assembly accuracy of Si-detectors in the modules and planes FwSD
E.V. Zubarev, O.G. Tarasov behalf on Forward Silicon Tracker team
10th Collaboration Meeting of the BM@N Experiment at the NICA Facility

## Slide 2
The accuracy of the assembly is based on:
On the accuracy of cutting Si-detectors using a diamond disc (through) - ±20 microns (measurement results slide # 4);
On the accuracy of assembly by the conductor of Si-detectors with a frame - ±30 μm (measurement results of the assembled modules, the inaccuracy of clause 1. is included in ±30 μm);
After assembling the plane, the true position of each Si-detector is measured on a video measuring microscope NVM II-5040D "NORGAU" with an accuracy of ±20 μm. Measurements are transmitted in the form of a file (table) to the physical group for analysis and reconstruction of "hits".
The modules are installed in two levels with overlap.
Measurement of the true position of each Si-detector relative to the coordinate system (X,Y,Z) of the outer reference pins of the plane
+
beam pipe position
reference pin
reference pin
Y
X
Si #1
Si #2
Si #1
Si #2
Assembly of half-planes
Assembling the module
Si #1
Si #2
module frame
126 mm
63 mm
vacuum
Ø50
conductor

## Slide 3: Measurements of the position of Si-detectors in modules and planes were carried out on a non-contact video measuring microscope "NORGAU" NVM II-5040D.
«NORGAU» NVM II-5040D | 
X- and Y-axis move range (mm) | 500 х 400
Z-axis move range (mm) | 250
Permissible absolute error of linear measurements along the X and Y axes (μm)* | ±(2.5+L/200)
Permissible absolute error of linear measurements along the Z-axis (μm)* | ±(2.5+L/100)
* L - measured length in mm | 
video measuring microscope
Operator's table
object to be measured (5 modules plane)
lens
subject table
no wire
Si #1
Si #2
detector reference mark
Topological signs on the detector can be used as reference points
object to be measured

## Slide 4
Measurement of cutting accuracy of Si detectors (mm)
sensor | a | b | c | d | e | f | g | h
#13в | 0,588 | 0,595 | 0,593 | 0,591 | 0,235 | 0,227 | 0,232 | 0,239
#13н | 0,598 | 0,603 | 0,587 | 0,581 | 0,244 | 0,247 | 0,238 | 0,236
#15/2016в | 0,595 | 0,596 | 0,592 | 0,591 | 0,234 | 0,236 | 0,233 | 0,236
#15/2016н | 0,595 | 0,594 | 0,591 | 0,595 | 0,233 | 0,233 | 0,235 | 0,232
#15в | 0,608 | 0,599 | 0,594 | 0,605 | 0,233 | 0,246 | 0,253 | 0,240
#15н | 0,603 | 0,606 | 0,588 | 0,587 | 0,234 | 0,233 | 0,251 | 0,258
#16/2017в | 0,600 | 0,595 | 0,589 | 0,591 | 0,232 | 0,230 | 0,234 | 0,234
#16/2017н | 0,602 | 0,601 | 0,581 | 0,587 | 0,242 | 0,244 | 0,225 | 0,244
#17/2016в | 0,595 | 0,593 | 0,595 | 0,597 | 0,231 | 0,232 | 0,236 | 0,240
#17/2016н | 0,595 | 0,597 | 0,593 | 0,593 | 0,234 | 0,236 | 0,230 | 0,230
#43в | 0,596 | 0,595 | 0,598 | 0,601 | 0,246 | 0,243 | 0,233 | 0,239
#43н | 0,607 | 0,603 | 0,596 | 0,592 | 0,253 | 0,249 | 0,236 | 0,238
#44в | 0,587 | 0,576 | 0,627 | 0,615 | 0,263 | 0,242 | 0,227 | 0,245
#44н | 0,584 | 0,579 | 0,612 | 0,609 | 0,259 | 0,256 | 0,220 | 0,222
#45в | 0,592 | 0,582 | 0,614 | 0,626 | 0,259 | 0,269 | 0,236 | 0,229
#45н | 0,585 | 0,576 | 0,609 | 0,622 | 0,246 | 0,261 | 0,232 | 0,218
The spread of the position of the edge of the cutting line relative to the reference crosses is ± 20 μm
595±20μm
235±20μm
detector reference mark
Pads for US-bonding
p+-strips
cutting line
Side of p+-strips
The position of the detector cut line, perpendicular to the strips
10μm
Position of the detector cut line, parallel to the strips
10μm
distance from the cutting line to detector reference mark along the X axis
number of measured sides
distance from the cutting line to detector reference mark along the Y axis
number of measured sides

## Slide 5: Measurement of the accuracy of the position of Si-detectors in the modules
Si #1
Si #2
module frame
126 mm
63 mm
Modul # | Хв (mm) | Ув (mm) | Хн (mm) | Ун (mm) | ϕв (deg.) | ϕн (deg.) | A (mm) | α (deg.)
#13 | 6.084 | 151.994 | 6.044 | 88.974 | 89.993 | 89.998 | 0.031 | 0.005
#15/2016 | 6.082 | 152.029 | 6.053 | 89.001 | 89.969 | 89.993 | -0.006 | 0.025
#15 | 6.094 | 152.003 | 6.046 | 88.978 | 89.974 | 89.988 | 0.019 | 0.014
#16/2017 | 6.079 | 152.031 | 6.047 | 88.987 | 89.973 | 89.994 | 0.002 | 0.021
#17/2016 | 6.112 | 152.032 | 6.038 | 88.980 | 89.960 | 89.997 | 0.031 | 0.037
#43 | 6.078 | 152.012 | 6.064 | 88.968 | 89.973 | 89.972 | -0.015 | -0.001
#44 | 6.122 | 152.026 | 6.101 | 88.972 | 89.936 | 89.954 | -0.028 | 0.018
#45 | 6.128 | 152.002 | 6.074 | 88.936 | 89.967 | 90.000 | 0.017 | 0.033
20 μm
shift between reference marks ±30 μm
alignment line
X-strip
Baseline through pins centers
ϕн
ϕв
ϕ = 89,98°
+0,02 °
- 0,04 °
side of inclined strips
side of inclined strips
Baseline through pins centers
The accuracy of the assembly is ± 30 microns according to the measurement results.

## Slide 6: The true position of Si-detectors in the half-plane (5 modules). (The measurement data is transmitted to the physical group).
detector location | X±0.02* (mm) | Y±0.02* (mm) | rotation in the plane 0XY (deg.) | Z±0.2** (mm) | Module Serial Number
Si-1a | 65.63 | 164.25 | 0.03 cw | 15.9 | #34
Si-1b | 65.64 | 101.26 | 0.01 cсw | 15.7 | 
Si-2a | 125.54 | 164.23 | 0.04 cсw | 27.5 | #16
Si-2b | 125.61 | 101.26 | 0.01 сw | 27.5 | 
Si-3a | 248.69 | 134.23 | 0.02 сw | 14.5 | #41
Si-3b | 248.60 | 71.23 | 0.03 сw | 14.2 | 
Si-4a | 308.50 | 164.25 | 0.01 сw | 26.1 | #24
Si-4b | 308.51 | 101.26 | 0 | 26.0 | 
Si-5a | 368.71 | 164.27 | 0.07 cw | 14.4 | #21
Si-5b | 368.61 | 101.27 | 0.03 cw | 14.2 | 
Position of Si-detectors in the half-plane
# 5/2 (06.07.22)
*-position of the origin point of the Si-detectors in the coordinate plane 0XY (tied to the outer reference pins).
**-the distance from the base surface of the base bar to the middle of the Si-detector surface farthest from it.
reference pin
reference pin
base bar
base surface
Si-detectors
Side of inclined strips (N+)
Side of parallel strips (P+)
#34
#16
#41
#24
#21

## Slide 7: Results of measurements of the location of half-planes in the BM@N-2022 (physical session - run Xe).
X= 244.49
Y= 140.28
Z= 190.76
X= 243.22
Y= -143.73
Z= 140.71
X= 245.74
Y= -175.67
Z= 226.68
X= 367.06
Y= -177.24
Z= 396.97
X= 369.33
Y= 177.75
Z= 451.16
X= -287.07
Y= 173.01
Z= 367.79
X= -228.26
Y= -177.19
Z= 229.44
X= -229.46
Y= 141.92
Z= 191.76
X= -230.78
Y= -143.06
Z= 141.75
X= -344.63
Y= 179.33
Z= 454.28
X= -346.93
Y= -174.45
Z= 406.22
Measurements 07.02.2023
SC Description:
- origin point (0,0,0) is the geometric center of the edge of the lower pole of the magnet SP - 41.
- the Z axis is directed along the beam axis;
- the Y axis is directed upwards perpendicular to the horizon;
Z
Y
X
X= 306.81
Y= 174.19
Z= 365.13
X= 306.77
Y= -175.32
Z= 314.48
X= 245.57
Y= 173.80
Z= 278.85
Data from the report of LLC "Industrial Measurements " “Measurement of the actual position of the detectors of the BM@N installation and their alignment to the design position relative to the beam – 2022.”
X= -287.15
Y= -176.48
Z= 312.86
X= -228.51
Y= 172.87
Z= 279.80
0,25°
0,74°
0,26°
0,16°
0,11°
0,33°
0,12°
0,13°
Top view
X
Z
#9/2
#9/1
#7/2
#7/1
#5/2
#5/1
#3/2
#3/1
Linking half-plane bases to geodetic measurement points

## Slide 8: True positions of Si-detectors on all planes by session BM@N-2022 (physical session -run Xe). The data is transferred to the physical group.
Позиция сенсора | X±0.02* (мм) | Y±0.02* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1 | 125.48 | 130.98 | 0.04 пр. час. | 27.9 | #3-0/22
Si-2 | 248.54 | 101.10 | 0.03 пр. час. | 14.0 | #2-0/22
Si-3 | 308.48 | 131.00 | 0 | 26.1 | #1-0/22
Положение Si-сенсоров в полуплоскости
# 3/1 (12.07.22)
Позиция сенсора | X±0.02* (мм) | Y±0.02* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1 | 125.56 | 130.86 | 0.02 пр. час. | 27.0 | #5-0/22
Si-2 | 248.52 | 101.13 | 0.04 пр. час. | 14.5 | #6-0/22
Si-3 | 308.50 | 131.12 | 0.01 по час. | 26.0 | #4-0/22
Положение Si-сенсоров в полуплоскости
# 3/2 (12.07.22)
Позиция сенсора | X±0.02* (мм) | Y±0.02* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1a | 65.51 | 164.29 | 0.04 пр. час. | 14.9 | #35
Si-1b | 65.66 | 101.19 | 0.06 пр. час. | 15.3 | 
Si-2a | 125.53 | 164.25 | 0.01 пo час. | 27.2 | #38
Si-2b | 125.57 | 101.18 | 0.01 пр. час. | 27.4 | 
Si-3a | 248.43 | 134.28 | 0.04 пo час. | 14.6 | #42
Si-3b | 248.47 | 71.27 | 0.01 пo час. | 14.3 | 
Si-4a | 308.52 | 164.23 | 0.03 пo час. | 25.9 | #27
Si-4b | 308.49 | 101.24 | 0.01 пo час. | 26.0 | 
Si-5a | 368.43 | 164.32 | 0.01 пр. час. | 14.3 | #32
Si-5b | 368.42 | 101.32 | 0.07 пр. час. | 14.2 | 
Положение Si-сенсоров в полуплоскости
# 5/1 (11.07.22)
Позиция сенсора | X±0.02* (мм) | Y±0.02* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1a | 65.63 | 164.25 | 0.03 пo час. | 15.9 | #34
Si-1b | 65.64 | 101.26 | 0.01 пр. час. | 15.7 | 
Si-2a | 125.54 | 164.23 | 0.04 пр. час. | 27.5 | #16
Si-2b | 125.61 | 101.26 | 0.01 пo час. | 27.5 | 
Si-3a | 248.69 | 134.23 | 0.02 пo час. | 14.5 | #41
Si-3b | 248.60 | 71.23 | 0.03 пo час. | 14.2 | 
Si-4a | 308.50 | 164.25 | 0.01 пo час. | 26.1 | #24
Si-4b | 308.51 | 101.26 | 0 | 26.0 | 
Si-5a | 368.71 | 164.27 | 0.07 пo час. | 14.4 | #21
Si-5b | 368.61 | 101.27 | 0.03 пo час. | 14.2 | 
Положение Si-сенсоров в полуплоскости
# 5/2 (06.07.22)
Позиция сенсора | X±0.02* (мм) | Y±0.02* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1a | 65.59 | 164.17 | 0.05 пр. час. | 27.7 | #14
Si-1b | 65.48 | 101.15 | 0.05 пр. час. | 27.6 | 
Si-2a | 125.53 | 164.26 | 0.06 пр. час. | 15.8 | #36
Si-2b | 125.57 | 101.28 | 0.05 пр. час. | 15.6 | 
Si-3a | 185.55 | 164.18 | 0.07 пр. час. | 27.6 | #9
Si-3b | 185.57 | 101.18 | 0.06 пр. час. | 27.5 | 
Si-4a | 308.56 | 134.28 | 0.03 по час. | 13.9 | #33
Si-4b | 308.47 | 71.31 | 0.03 пр. час. | 13.9 | 
Si-5a | 368.61 | 164.27 | 0.02 по час. | 25.4 | #30
Si-5b | 368.53 | 101.27 | 0 | 25.9 | 
Si-6a | 428.63 | 164.26 | 0.02 по час. | 14.2 | #15
Si-6b | 428.55 | 101.27 | 0.01 по час. | 14.0 | 
Si-7a | 488.77 | 164.20 | 0.08 по час. | 25.6 | #29
Si-7b | 488.63 | 101.21 | 0.08 по час. | 26.0 | 
Положение Si-сенсоров в полуплоскости # 7/1 (17.03.22)
Позиция сенсора | X±0.02* (мм) | Y±0.02* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1a | 65.53 | 164.30 | 0 | 26.9 | #7
Si-1b | 65.53 | 101.32 | 0.02 по час. | 27.2 | 
Si-2a | 125.36 | 164.20 | 0.07 пр. час. | 15.7 | #22
Si-2b | 125.43 | 101.20 | 0.03 пр. час. | 15.5 | 
Si-3a | 185.55 | 164.31 | 0.01 пр. час. | 27.3 | #2-21
Si-3b | 185.56 | 101.29 | 0.01 по час. | 27.4 | 
Si-4a | 308.44 | 134.11 | 0.02 по час. | 14.4 | #6-21
Si-4b | 308.41 | 71.12 | 0 | 14.1 | 
Si-5a | 368.50 | 164.27 | 0.03 по час. | 27.1 | #4-21
Si-5b | 368.47 | 101.27 | 0.02 по час. | 26.2 | 
Si-6a | 428.30 | 164.27 | 0.03 пр. час. | 14.0 | #31
Si-6b | 428.33 | 101.29 | 0.04 пр. час. | 13.9 | 
Si-7a | 488.47 | 164.24 | 0.01 по час. | 26.0 | #1-20
Si-7b | 488.46 | 101.22 | 0.04 по час. | 25.9 | 
Положение Si-сенсоров в полуплоскости # 7/2 (17.03.22)
*-положение точки начала координат Si-сенсора в координатной плоскости 0XY (привязана к наружным базовым отверстиям).
**-взгляд на лист.
***-расстояние от базовой поверхности базовой рейки до середины дальней от неё поверхности Si-сенсора.
Позиция сенсора | X±0.03* (мм) | Y±0.05* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1a | 65.56 | 164.26 | 0.03 пр. час. | 15.4 | #12
Si-1b | 65.50 | 101.26 | 0.05 пр. час. | 15.5 | 
Si-2a | 125.57 | 164.19 | 0.04 пр. час. | 26.7 | #5-21
Si-2b | 125.52 | 101.27 | 0.03 пр. час. | 27.3 | 
Si-3a | 185.64 | 164.34 | 0 | 15.8 | #25
Si-3b | 185.54 | 101.35 | 0.01 по час. | 15.7 | 
Si-4a | 245.49 | 164.30 | 0.07 пр. час. | 27.4 | #18
Si-4b | 245.47 | 101.32 | 0.06 пр. час. | 27.5 | 
Si-5a | 368.53 | 134.35 | 0.02 по час. | 14.7 | #31
Si-5b | 368.45 | 71.37 | 0.01 пр. час. | 14.4 | 
Si-6a | 428.51 | 164.30 | 0 | 25.9 | #40
Si-6b | 428.45 | 101.30 | 0.01 пр. час. | 26.1 | 
Si-7a | 488.70 | 164.33 | 0.03 по час. | 14.5 | #23
Si-7b | 488.56 | 101.33 | 0.05 по час. | 14.4 | 
Si-8a | 548.54 | 164.33 | 0.01 пр. час. | 25.6 | #26
Si-8b | 548.46 | 101.34 | 0.01 пр. час. | 25.9 | 
Si-9a | 608.68 | 164.30 | 0.05 по час. | 14.6 | #11
Si-9b | 608.55 | 101.28 | 0.02 по час. | 14.4 | 
Положение Si-сенсоров в полуплоскости # 9/1 (31.03.22)
Положение Si-сенсоров в полуплоскости # 9/1 (31.03.22)
Позиция сенсора | X±0.03* (мм) | Y±0.05* (мм) | Разворот в плоскости 0XY (град.)** | Z±0.2*** (мм) | Серийный номер модуля
Si-1a | 65.61 | 164.27 | 0.01 пр. час. | 15.8 | #39
Si-1b | 65.61 | 101.24 | 0 | 15.7 | 
Si-2a | 125.49 | 164.25 | 0.02 пр. час. | 27.6 | #8 #9-2017
Si-2b | 125.54 | 101.22 | 0.01 по час. | 27.6 | 
Si-3a | 185.46 | 164.18 | 0.06 пр. час. | 15.9 | #28
Si-3b | 185.52 | 101.17 | 0.04 пр. час. | 15.7 | 
Si-4a | 245.55 | 164.26 | 0.03 пр. час. | 27.5 | #20
Si-4b | 245.58 | 101.25 | 0.02 пр. час. | 27.6 | 
Si-5a | 368.53 | 134.17 | 0.02 по час. | 14.3 | #10
Si-5b | 368.49 | 71.15 | 0.01 пр. час. | 14.1 | 
Si-6a | 428.51 | 164.20 | 0.02 по час. | 26.1 | #19
Si-6b | 428.47 | 101.20 | 0.01 пр. час. | 26.0 | 
Si-7a | 488.67 | 164.13 | 0.09 по час. | 14.9 | #3-21 #10-2017
Si-7b | 488.57 | 101.13 | 0.06 по час. | 14.3 | 
Si-8a | 548.57 | 164.14 | 0.02 по час. | 26.1 | #13
Si-8b | 548.52 | 101.12 | 0.02 по час. | 26.1 | 
Si-9a | 608.67 | 164.14 | 0.05 по час. | 14.1 | #17
Si-9b | 608.57 | 101.14 | 0.04 по час. | 14.0 | 
The angles of rotation of the detectors in half-planes. View from the measurement side
Si-detectors rotations (deg.)
number of detectors

## Slide 9
Conclusions:
The correct method of assembling Si-detectors into a module based on the positioning of Si-detectors by the cutting line on the conductor has been chosen;
Measuring the position of each Si-detector after assembling the planes compensates for inaccuracies in intermediate assembly operations;
In preparation for the next session of the BM@N experiment, the installation of planes should be carried out under the supervision of surveyors;
By the next session of the BM@N experiment, make changes to the design of the planes to install additional geodetic marks

## Slide 10
Спасибо за внимание!