## Slide 1: Stability of 136Ba peaks and FWHM from Oct and Nov 2021 data sets
-Dhanurdhar bajpai

## Slide 2: Outline
Review of procedure for time stability analysis and calibration correction
Variation of Peaks between Oct & nov data
  REGE detectors
  BEGE detectors
  COAX detectors
Variation of FWHM between Oct & nov data
  REGE detectors
  BEGE detectors
  COAX detectors
Summary

## Slide 3: Procedure for time stability analysis
Time stability analysis is needed to understand the detector performance during the beamtime
This includes shift in muonic X-ray peak positions and the peak FWHMs
136Ba data was divided into sets of 6 hours and all files within the 6 hour period were combined
4 peaks (1227.0 and 1283.7 KeV in low energy and 3922.9 and 2988.8 keV in high energy region) were selected and the mean/ sigma/ residuals were found using the GERDA fit function
1460.83 keV 40K peak was used as a cross check of our detector stability over the entire (136Ba+76Se) measurement
This talk only focusses on 136Ba time stability

## Slide 4: Example of Ba and K peak stability over time

## Slide 5: Procedure for calibration correction
136Ba had two sets of data, Oct & Nov using calibrations from Oct 20 and Nov 7 respectively
Oct data with Oct 20 calibrations had worse peak mean positions and poor residuals

## Slide 6: Procedure for calibration correction
So, we calculated the peak positions again by multiplying it with a correction factor for each detector as follows:
  Factor =
This gave approximate peak positions without generating new tier3 files

## Slide 7: Cross-check calibration correction with 40K

## Slide 8: Variation of Peaks between Oct & nov data

## Slide 9: Channel 0 (REge) peaks and residuals
Shift
Stable
Poor residuals

## Slide 10: Channel 4 (REge) peaks and residuals
Stable
Stable
Poor residuals

## Slide 11: Channel 1 (bege) peaks and residuals
Stable
Shift
Poor residuals

## Slide 12: Channel 7 (Coax) peaks and residuals
Stable
Shift
Poor residuals

## Slide 13: Shift/RMS for all detectors
|Shift|= Difference in means
FWHM is calculated from fit parameters of at that energy from calibration data
Excluded first 1.5 days of data

## Slide 14: Shift/RMS for all detectors
Crude estimates of shift and RMS variation between Oct and Nov data
“Shift” is subject to threshold we set for the analysis (say 1.5 keV ?)
Peaks are “stable/unstable” if shifts are smaller/greater than this threshold
We also look at the FWHM of the peak to decide if we combine those data sets*
RMS can be a good estimate of peak stability within individual periods
*The FWHM from physics data is larger than calculated one (later slides)

## Slide 15: Observation by looking at the peaks
Summary of peak stability of all detectors separately in low/high energy region
Magnitude of shift in peaks between end of Oct and beginning of Nov

## Slide 16: Variation of FWHM between Oct & nov data

## Slide 17
Channel 0
Channel 3
Channel 2
Channel 4
REGE DETECTORS FWHM
Eyeball observation: Behavior of FWHM is similar in given energy range/ detector type

## Slide 18
Channel 1
Channel 5
BEGE DETECTORS FWHM

## Slide 19
Channel 6
Channel 7
COAX DETECTORS FWHM
A jump in FWHM is seen from late Oct to initial Nov data but it lowers after the first few hours (~1-1.5 days)

## Slide 20
Study the % change in FWHM of detectors by their type
% change from the calculated mean of FWHM from Oct & Nov data sets
Change in FWHM is within 5-15% for REGE detectors
Channels 2,3 and 4 have 0-2% change in FWHM from Oct to Nov at higher energies
Not true for Channel 0; has substantially high FWHM change at higher energy
Could be since this channel is the worst resolution detector compared to other channels (result from 152Eu FWHM vs Energy curve and 1.3 MeV 60Co peak)
< 2% change in FWHM for BEGE detectors
5-15% change in FWHM for COAX detectors

## Slide 21
REGE DETECTORS FWHM
Higher energies have higher FWHM of the peaks in both data sets
Change in FWHM is within 5-15% for REGE detectors
Channels 2,3 and 4 have 0-2% change in FWHM from Oct to Nov at higher energies
Not true for Channel 0; has substantially high FWHM change at higher energy
Could be since this channel is the worst resolution detector compared to other channels (result from 152Eu FWHM vs Energy curve and 1.3 MeV 60Co peak)

## Slide 22: Summary
We need to generate new tier3 files applying October 28 calibrations to first set of 136Ba data and all the plots shown have the Oct 28 calibrations
The residuals are bad for higher energies as our calibration is only up to 1.4 MeV
We need to set a threshold for shift to call them stable/unstable, additionally look at the FWHM to decide if we combine the data sets or not
If there is a shift between the Oct and Nov data sets, for analysis, we will not combine the two periods, rather only runs within the same stable period since we cannot apply any time correction
The worse FHWM in Nov data set is mainly because of the change in dynamic range
There is a change in FWHM of peaks, mostly depending on detector type, bad for REGE and COAX detectors, quite stable for BEGE ones
Another separate report on time stability is in progress which documents all the plots and information and will be soon shared

## Slide 23: Backup slides

## Slide 24: Channel 2 (REge) peaks and residuals

## Slide 25: Channel 3 (REge) peaks and residuals

## Slide 26: Channel 5 (bege) peaks and residuals

## Slide 27: Channel 6 (Coax) peaks and residuals