## Slide 1: Comparison of two methods for processing signals from neutron counters based on a scintillator 6LiF/ZnS(Ag) and optical fibers
Podlesnyy M. M., Bogdzel А. А., Milkov V. M.
Alushta-2023

## Slide 2: Outline
Introduction
Experimental setup
Results
Conclusion

## Slide 3: Detector efficiency
Efftotal= Effconv * Effsignal
1 signal
Mean signal (500 avg.)

## Slide 4: Comparison of 2 methods
Pulse shape discrimination method (PSD)
Digitiser
N6730,
Caen
Timing filter
MPD-32
PSD = Q2/Q1
t > tveto
tdead time

## Slide 5: Formulation of the experiment
Goal – comparison of 2 methods at different neutron flux.
Input parameters should be the same for both methods.
Criteria for false event identification should be the same for both methods.
Subject – neutron count rate, percentage of false counts.
Variables: Dead time, Analog shaping time.

## Slide 6: False event criteria
Recording shape of signals (too long and difficult)
Δt between consecutive signals (MPD-32)
Comparison events count rate with «non-paralyzable dead time curve»: N = n*t/(1+n𝜏)
The main problem is to guarantee right identification of neutron events

## Slide 7: Experimental setup
2 neutron sources (252Cf incapsulated and neutron generator)
2 neutron counters (50 cm2 and 1400 cm2)
PMT signal supplies on amplifier (х50) with implemented RC integrated chain (45, 130, 300 ns), then doubled and supplies on MPD-32 and digitiser N6730, Caen
ASTRA-M detector element (1400 cm2)
Neutron generator (~ 2*107 n/s in 4pi)

## Slide 8: Measurements with 252Cf and a test counter (~100 n/s)
Tshaping = 45 ns
t of measurement = 1000 s
Threshold = 21 mV
MPD-32
tveto = 96 ns
Дигитайзер
Gates 100 ns и 1000 ns

## Slide 9: Measurements with 252Cf and a test counter (~100 n/s)
% of false counts with low sample rate <6% for MPD-32
PSD method work properly for d.t.≥ 5 mks
Amplitude spectrum
% of false counts
Dead time
Additional amplitude filter

## Slide 10: Measurements with a neutron generator and ASTRA-M counter
filtered events trigger dead time lasting ~20% of the entire measurement
Shaping time 45 ns
t of measurement = 100 s
Threshold = 21 mV

## Slide 11: Measurements with a neutron generator and ASTRA-M counter
an increase in shaping time led to a decrease in the efficiency by 33%
Shaping time 300 ns
t of measurement = 100 s
Threshold = 21 mV

## Slide 12: MPD-32 counting curves with different thresholds
Shaping time 45 ns
MPD-32 counts ~9kHz properly, then curve falls (66% loss)
Digitiser counts neutrons properly 2 times more

## Slide 13: Conclusion
PSD method work properly at high neutron flux.
PSD method with low threshold requires minimum dead time 5 ≥ mks (~12% loss)
Real efficiency loss due to signal processing is ~20-33%.
MPD-32 is very sensitive to setup and has a high losses at high neutron flux (~66%). Why?

## Slide 14: Thanks for your attention!

## Slide 15: Discussions, discussions, discussions…

## Slide 16: Measurements with neutron generator and a small counter
Shaping time 45 ns

## Slide 17: Time filter method

## Slide 18: PSD method