## Slide 1
Cryogenic system of SPD
Speaker: PhD Dr. Nikiforov Dmitry

## Slide 2
Outline
Cryogenic system
Helium system
3. Nitrogen system
4. Steps of creation
5. Conclusions

## Slide 3
Cryogenic system
Helium pipelines
N2, Ar, CO2 pipelines
Cryogenic storage tanks
Helium liquefier

## Slide 4
Helium system

## Slide 5
 | Operating parameters | Unit
1 | Cooling capacity (for 4.5 K) | 100 - 130 l/h
 | Cooling capacity (for 50 K) | 150 W
2 | Temperature of outlet flow from the SPD | 4.3 K (1.05 bar)
3 | Temperature of inlet flow from the SPD | 4.5 К (1.15 bar)
4 | Hydraulic resistance of the SC coil | 0.1 bar
5 | Cold weight | 4000 kg
6 | Maximum pressure in coil | 5 MPa
7 | Heat load | 60 – 80 W
8 | Equipment Requirement | Maximum reliability, energy efficiency, compactness, automatic mode
9 | Interval of repair/regulatory work of the plant | Not more than once a year
Helium system

## Slide 6
Service position
Working position
Helium system
Power supply
Control cryostat
Liquefier

## Slide 7
Helium system
First option

## Slide 8
Nitrogen system
First option

## Slide 9
Service position
Working position
Helium system

## Slide 10
Nitrogen system
Second option

## Slide 11
Detectors/Gas | N2 | Ar | CO2 | C4H10
RS | 0,4 t | 0,9 t | 0,4 t | -
ST | - | 4,5 t | 2,1 t | -
MM | - | 0,029 t | - | 0,0045 t
Cryo plant | ≈ 15 t | - | - | -
Total | 15,4 t | 5,429 t | 2,5 t | 0,0045 t
Gas consumption for SPD detectors during 1 month of operation
Nitrogen system

## Slide 12
Steps of creation |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
SPD | 2022 |  |  |  | 2023 |  |  |  | 2024 |  |  |  | 2025 |  |  |  | 2026 |  |  |  | 2027 |  |  | 
 | I | II | III | IV | I | II | III | IV | I | II | III | IV | I | II | III | IV | I | II | III | IV | I | II | III | IV
Cryogenic plant |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Helium pipeline |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Nitrogen system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Vacuum system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Pneumatic system |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
Cryogenic system control |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  | Technical task |  |  |  |  | Contract execution |  |  |  |  |  | Assembly & Test |  |  |  |  | Commissioning |  |  |  |  |  | 

## Slide 13
Conclusions:
The manufacturer of the SPD magnet and the type of cryogenic plant were determined.
The location of cryogenic tanks, their volumes and gas consumption for SPD were determined.
The maximum cooling capacity of the cryogenic plant is calculated.
The development of technical specifications for a cryogenic plant and a platform for cryogenic tanks is in an active phase.
The work is being carried out in accordance with the plan.

## Slide 14
Thank you for your attention!