## Slide 1
NDL SiPm Series EQR15 11-6060D-S
1) Test results with 480 nm LED
2) Cosmic test www.ndl-sipm.net/PDF/Datasheet-EQR15.pdf
12/19/2023
1
Oleg Gavrishuk, JINR, Dubna , Russia
NDL (Novel Device Laboratory, Beijing) http://www.ndl-sipm.net/indexeng.html

## Slide 2
NDL SiPm Series EQR15 11-6060D-S www.ndl-sipm.net/PDF/Datasheet-EQR15.pdf
For a conventional SiPM, the quenching resistors are usually fabricated on the surface, and used to connect all APD cells to trace metal lines. In contrast, NDL SiPM employs intrinsic epitaxial layer as the quenching resistors (EQR), and uses a continuous silicon cap layer as an anode to connect all the APD cells. As a result, the device has more compact structure and simpler fabrication technology, allows larger micro cell density (larger dynamic range) while retaining high photon detection efficiency (PDE).
Для обычного SiPM гасящие резисторы обычно изготавливаются на поверхности и используются для соединения всех ячеек APD с металлическими линиями. Напротив, NDL SiPM использует собственный эпитаксиальный слой в качестве гасящих резисторов (EQR) и использует непрерывный слой кремния в качестве анода для соединения всех ячеек APD. В результате устройство имеет более компактную структуру и более простую технологию изготовления, позволяет увеличить плотность микроячеек (больший динамический диапазон) при сохранении высокой эффективности детектирования фотонов (PDE).
NDL (Novel Device Laboratory, Beijing) http://www.ndl-sipm.net/indexeng.html
12/19/2023
2
Oleg Gavrishuk, JINR, Dubna , Russia

## Slide 3
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
3

## Slide 4
Photon Detection Efficient (PDE) correspond to WLS Emission spectra of Y11.
PDE close to flat maximum about 45% at 6-7 Overvoltage.
Notes: 4

## Slide 5
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
5
EQR11-15-6060-S
Front – 10 ns
Length – 20 ns
Pulse shape of SiPm with 15 µ pitch and 6x6 mm2 size – frm 10 ns LED
HAMAMATSU S14160-50
Front – 21 ns
Length – 89 ns

## Slide 6: Wave form and digitizing

## Slide 7: LED amplitudes to caclucte Number Pixel Fired = {Amean/sigma}**2

## Slide 8
SiPm test Results
12/19/2023
8
Oleg Gavrishuk, JINR, Dubna , Russia
Dark Current measurements
Gain vs Overvoltage: dA/dU
Gain vs Temperature: dA/dt
PDE vs Overvoltage
Operation Volt. how to select it ?
Temperature stability studies
Chart: Temperatute, 0C
| 0 | 10 | 20 | 30
38 V | 357.75 | 420.075 | 574.5 | 870.25
36 V | 116.46 | 185.16 | 293.2 | 473.2

## Slide 9
EQR15-60 Dark Current vs Bias voltage and its temperature dependences around operate Voltage
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
9
Chart: Temperature, 0C
| 0 | 10 | 20 | 30
38 V | 357.75 | 420.075 | 574.5 | 870.25
36 V | 116.46 | 185.16 | 293.2 | 473.2
Dark Current vs Operation Bias (36-37 V) at room temperature (20 0C)
is equal to 300-400 nA. Its are corresponded to the factory data and is similar HAMAMATSU too.

## Slide 10
Gain vs Overvoltage allow obtain: slope=dA/dU vs U
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
10
T=-10 0C
T=+30 0C
Gain vs Temperature allow obtain:
slope=dA/dT vs T
U=38V
U=32V
Kt=dU/dT=dA/dT/dA/dU
The thermal stabilization coefficient (Kt) was defined as the ratio of the slope dA/dT to the slope dA/dU depending on the applied bias.
dA/dU
dA/dT

## Slide 11
Break Point (Bp) was defined as extrapolation point of dA/dU to zero. Take assumption his linear behavior from Temperature we find that Bp=32 V at 20 0C.
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
11
Kt=dU/dT=dA/dt/dA/dU fom previous slides we find that dependence from Overvoltage has linear behavior and equal to ~50 mv/0C at Operation point 5.5 V.
Operate point

## Slide 12
Temperature sensitivity vs dU
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
12
dA/dT vs dU
PDE vs dU
Photon Detection Efficient vs dU
Operation point
Operation point
Operation point was found at dU=6.5 V take in account that:
dA/dT and PDE – has Plato in dU.
dA/dT has minimal value ~ 1%/0C , PDE close to maximal value ~ 32%.

## Slide 13: Kt=dA/dT / dA/dU

## Slide 14: dA/dT % vs Overvotage

## Slide 15: Test results with cosmic particles
4 Module 3x3 cell wit sixes 40x40 mm – MPD option 2012
Cosmic beam
Event selection = only One cell – one hit events
LED monitoring – 1 Hz permanently
Kt = 20,30,40,50 mv/Grag for different modules
Mip resolutin = 9.7 %
Long time stabilyti at Kt=50 mv/Grad

## Slide 16
Test results with cosmic particles
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
16
Setup of 4 modules. Each module consist from 9 cells of 4x4 cm2. Totally tested 36 cells.
Cosmic Rays
Single Ecal module shown in assembling stage.
It is visible 9 cells as 3x3 matrix with WLS fibers (16 per cell). Y11(200) diameter 1.0 mm was used.
Sampling:
1.5 mm Scintillator
0.3 mm Lead
200 layers
Scintillator composition:
Polysterene
1.5% Paterphenyle
0.04% POPOP

## Slide 17
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
17
MIP spectra from 36 Cells. Top view shown on picture Above. One hit/event – applied selection criteria during analysis.
Calibration coefficients were found and normalized to 240 MeV.

## Slide 18
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
18
MIP spectra from 36 Cells as Total Sum take in account the Calibration coefficients normalized to 240 MeV. These Energy resolution corresponded to MC

## Slide 19
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
19
Test of long time Stability was done with different Kt for individual module.
The temperature variation from 20 to 30 0C per day was applied.
Kt=20 mV/0C
Kt=30 mV/0C
Kt=40 mV/0C
Kt=50 mV/0C

## Slide 20
12/19/2023
Oleg Gavrishuk, JINR, Dubna , Russia
20
Kt=50 mV/0C
Long time stability for 4 modules shown +/-0.5% LED signal variation during 45 hours at Kt=50 mv/0C.
Profile Sum total of MIP signal shown ~1% variation during 45 hours.
This variation is not sensitive for MIP spectra with 9% resolution.

## Slide 21: Matrix form for new scintillator production (40x40x1.5 mm3) should be ready in end of 2023
Matrix firm is special setup for molding by pressure technology
It allow produce scintillator from granulated polystyrene with doppands
4 set form produce 4 scintillator plate per 1 minute

## Slide 22: Plan for 2024: End Cup for SPD with new scintillator to be produced in February-March of2024 256 cells of 40x40 mm2 = 64 modules of 4 cells=2x2
This Figure shows in red 64 modules, consisting of 4 cells each. The weight of this assembly is 597 kg. This will require 130 kg of polystyrene, 465 kg of lead, as well as additives: 1.95 kg of P-terphenyl and 65 g. POPOP, and 2000 meters WLS fiber type Y-11.
It is 1/20 part of End Cup and
taken time of 36 Days to prepared 51200 Stint. plates.
To read this setup, we need four ADC64 - 64-channel amplitude encoders, as well as 16 boards of 16-channel amplifiers and bias voltage regulators.