## Slide 1: Algorithm for Material Analysis by the K-edges for Photon-Counting Computed Tomography
28th International Scientific Conference of
Young Scientists and Specialists
(AYSS-2024)
Sotenskii Rostislav
JINR, Dubna
sot.rostislav@gmail.com
29 October 2024, Dubna
JINR

## Slide 2
2/17
Multi-energy tomography
Sotenskii R./AYSS-2024/29.10.24
Multi-energy CT
where - mass attenuation
Anderson, Nigel G. et al. “Spectroscopic (multi-energy) CT distinguishes iodine and barium contrast material in MICE.” European Radiology 20 (2010): 2126-2134.
Bateman, Christopher & Rajendran, Kishore & Ruiter, Niels & Butler, Anthony & Butler, Philip & Renaud, Peter. (2015). The Hidden K-edge Signal in K-edge Imaging.

## Slide 3
3/17
Medipix detectors
Medipix series detectors are hybrid semiconductor pixel detectors;
Developed by Medipix collaboration (https://medipix.web.cern.ch/);
Consist of a semiconductor sensor and a readout integrated circuit;
Photons are detected by their transferring energy to electrons. The appeared free electrons move to the pixel contact pads, causing a signal.
The signal is digitized and compared with the threshold in a pixel. Pixels operate independent.
Sotenskii R./AYSS-2024/29.10.24

## Slide 4
4/17
Experimental Microtomograph “Kalan”
Widepix detector:
15 Medipix3RX in one row
256х3840 pixels
Size of pixel - 55x55 mkm
Si sensor
Several of our goals on this installation:
• Separate materials that differ in composition. In particular, select a contrast agent that has an absorption edge in the working energy range;
• Determine the main element of the contrast agent;
• Determine the concentration of the contrast agent.
In this case, it is necessary to use as little number energies as possible to reduce the scanning time.
Sotenskii R./AYSS-2024/29.10.24

## Slide 5
5/17
2D phantom for studying the "spectrum" obtained from the detector
Flat Field Correction:
FFC = F(object) / F(flat field)
FFC
Frame received from the detector in the form of a 2D histogram.
RoI – Region of Interest
FFC pixel value
Y, pixel
X, pixel
Sotenskii R./AYSS-2024/29.10.24

## Slide 6
6/17
Study of the "spectrum" obtained from the detector
In samples with a low concentration of Lanthanum the K-edges are almost invisible. The character of their spectra is the same as that of water, ethanol, and a iodine solution with low concentration. It is difficult to distinguish them.
The spectrum is the average pixel values ​​in the RoI after the FFC, depending on the threshold.
Sotenskii R./AYSS-2024/29.10.24

## Slide 7
7/17
The criterion for the identification of contrast agents
We define:
Conditions for the fulfillment of the criterion;
Contrast agent concentration.
Spectrum after normalization to trends.
Sotenskii R./AYSS-2024/29.10.24
a, b, c, d, m – energy thresholds for target CA

## Slide 8
8/17
The criterion for the identification of contrast agents
Conditions for fulfilling the criterion, written in general form through the primary spectrum :
Parameters a, b, c, d, m – equivalent to energy; differs for each type of contrast agent.
где l1, l2, l3 – solution of a system of equations describing a linear trend(1);
k1, k2, k3 – solution of a system of equations describing a parabolic trend(2)
(1)
(2)
Contrast agent concentration:
Sotenskii R./AYSS-2024/29.10.24

## Slide 9
9/17
Results of applying the criterion in 2D-CT (on the example of La)
FFC pixel value
Criterion value
Sotenskii R./AYSS-2024/29.10.24
The ROI for each sample was not determined in advance

## Slide 10
10/17
Results of applying the criterion in 3D-CT
Phantom «Eva»
Ø 70 mm
Reconstructed slice of the phantom
Sotenskii R./AYSS-2024/29.10.24

## Slide 11
11/17
Result of applying the criterion on a 3D phantom (La)
Phantom study with water solutions of lanthanum (La(NO3)3·6H2O) and additional samples (water, iodine, calcium, bone, paraffin)
Concentration of La calculated by the criterion
Fulfilment of the conditions of La.
Green – samples that have satisfied the conditions of the criterion.
Yellow – samples that have not satisfied the conditions of the criterion.
Slice of the phantom
Sotenskii R./AYSS-2024/29.10.24

## Slide 12
12/17
Result of applying the criterion on a 3D phantom (Gd)
Slice of the phantom
Concentration of Gd calculated by the criterion
Phantom study with water solutions of gadolinium (C14H20GdN3O10) and additional samples (water, iodine, calcium, bone, La solution)
The criterion was applied in each voxel and without selecting regions of interest
Sotenskii R./AYSS-2024/29.10.24

## Slide 13
13/17
Result of applying the criterion on a 3D phantom (Nd)
Slice of the phantom
Concentration of Nd calculated by the criterion
Phantom study with water solutions of neodinium (Nd(NO3)3·6H2O) and additional samples (water, iodine, calcium, bone, paraffin)
The criterion was applied in each voxel and without selecting regions of interest
Sotenskii R./AYSS-2024/29.10.24

## Slide 14
14/17
Result of applying the criterion on a 3D phantom
Dependence of the concentration calculated by the criterion on the actual concentration of contrast agents (La, Nd, Gd)
The values ​​at the zero point are taken to be the values ​​calculated for a sample with water, since the contrast agents under consideration are aqueous solutions.
Sotenskii R./AYSS-2024/29.10.24

## Slide 15
15/17
Separation of several contrast agents in one object
Phantom composition:
La-40, Nd-40, Gd-74 – samples of aqueous solutions;
*La-80, *Nd-80, *Gd-80 – samples based on graphite flakes;
H2O, Bone, Air, Plex – additional samples.
Slice of phantom
Dependence of gamma photon absorption on energy (detector threshold) for specific regions of interest
12 energy thresholds were used to identify and separate multiple contrast agents simultaneously
Sotenskii R./AYSS-2024/29.10.24

## Slide 16
16/17
Separation of several contrast agents in one object at the same time
Phantom composition:
La-40, Nd-40, Gd-74 – samples of aqueous solutions;
*La-80, *Nd-80, *Gd-80 – samples based on graphite flakes;
H2O, Bone, Air, Plex – additional samples.
Slice of phantom
3D CT images of phantom with CAs.
La – red
Nd – blue
Gd – green
Sotenskii R./AYSS-2024/29.10.24

## Slide 17
17/17
Conclusions
• The proposed method makes it possible to separate the contrast agent from other materials and classifying them;
• The developed criterion can be used to estimate the concentration of the contrast agent;
• The criterion was developed for lanthanum contrast agents, but it can be used for other contrast agents based on high-Z elements;
• This criterion can be used in 2D-CT and 3D-CT;
• 5 energies are required to separate one contrast agent;
• Multiple contrast agents can be separated at the same time.
Other methods for identifying and classifying contrast agents are also being tested. For example – machine learning.
This work has been supported by the Russian Science Foundation under grant № 22-15-00072.
Sotenskii R./AYSS-2024/29.10.24

## Slide 18: Thank you!

## Slide 19
7/19
Study of the "spectrum" obtained from the detector
Sotenskii R./AYSS-2024/29.10.24

## Slide 20
8/19
The criterion for the identification of contrast agents
First step: Normalization to a linear trend.
The first condition for fulfilling the criterion:
0
,
Contrast agent concentration:
Сcr = 1000( F1(с) – F1(d) )
where F1 – spectrum after normalization to a linear trend.
Sotenskii R./AYSS-2024/29.10.24

## Slide 21
9/19
The criterion for the identification of contrast agents
Second step: Normalization to a Parabolic trend.
Sotenskii R./AYSS-2024/29.10.24

## Slide 22
16/19
Разделение нескольких контрастных веществ в 3D-КТ
Срез фантома
Энергия К-края (thl) в областях интереса.
Красный цвет – положение K-края не найдено.

## Slide 23
17/19
Разделение нескольких контрастных веществ в 3D-КТ
RoI | Iodine | La | Nd | Gd
K-edge, thl | 85 | 99 | 110 | 121
K-edge, keV | 33.2 | 38.9 | 43.6 | 50.2
Зависимость энергии К-края, определенной по критерию, от экспериментальной энергии К-края

## Slide 24
The criterion for the identification of contrast agents
The first conditions for the fulfillment of the criterion:
F1(c) > F1(d)
F1(c) > 0
c
d
c
d
m
m

## Slide 25
Results of applying the criterion in 3D-CT (on the example of La)
The fulfillment of the conditions for the content of lanthanum in the RoI.
Green color - samples that meet the condition of the criterion.
Yellow color - samples that do not meet the criterion.
CT image of slice of the phantom
Investigated phantom with water solutions of La(NO3)3·6H2O and also the standard samples (water, iodine, paraffin, bone and plex)
The estimation of molar concentration of La calculated by the criterion

## Slide 26
Results of applying the criterion in 3D-CT (on the example of La)
RoI | La-40 | La-20 | La-10 | La-5 | Paraffin | Iodine | Ca-40 | H20 | Bone | Plex | Air
Criterion | Yes | Yes | Yes | Yes | No | No | No | No | No | No | No
Value | 281.5 | 145.7 | 46.8 | 27.7 | 5.6 | 725.9 | -48.2 | -13.6 | -157.8 | -22.9 | -7.6
Error | 39.3 | 37.9 | 37.2 | 37.6 | 37.6 | 55.6 | 38.3 | 37.0 | 42.5 | 13.9 | 11.9

## Slide 27
Results of applying the criterion in 3D-CT (on the example of Nd)
RoI | Nd-80 | Nd-40 | Nd-20 | Nd-10 | Paraffin | Iodine | Ca-40 | H20 | Bone | Plex | Air
Criterion | Yes | Yes | Yes | Yes | No | No | No | No | No | No | No
Value | 189.0 | 98.9 | 83.3 | 78.0 | 61.4 | -124.9 | 48.8 | 33.4 | 40.9 | -1.5 | -1.6
Error | 20.5 | 20.2 | 19.6 | 19.9 | 20.2 | 24.4 | 20.6 | 20.2 | 32.3 | 13.7 | 11.6

## Slide 28
Results of applying the criterion in 3D-CT
Slice of the phantom
Положение К-линии поглощения (thl)
The concentration of lanthanum calculated by the criterion.

## Slide 29
Separate several different CAs in 3D-CT
3D CT images of phantom with CAs and color division of La (red),
Nd (blue), and Gd (green)
Evaluation of the molar concentration of the CAs

## Slide 31
c
d
Range of
K-edge (La)
a
b

## Slide 32
a
b
m
c
d

## Slide 33
c
d
a
b

## Slide 34
c
d
a
b
m

## Slide 35
c
d
a
b
m