## Slide 1: Forward Silicon Tracker (FSD) status
Oleg Tarasov on behalf of Forward Silicon Tracker team
12th Collaboration Meeting of the BM@N Experiment at the NICA Facility
May 13 – 17, 2024

## Slide 2
Content of the report:
Diagnostics and elimination of defects in FSD-planes.
Coordinates of modules (Si-detectors) after disassembly-assembly of FSD-planes.
Refinment of FSD-planes positioning device in the BM@N experemental hall.
Conclusion.

## Slide 3: FSD: before replacing (Run 8305, 30.01.2023)
Summary:
10 dead zones, 3 were discovered before the run and 7 occurred during the run
#14
#9
#19
#17
#12
#5-21
#39
#23
Hits on station 1
6 modules, DSSD 63×93 mm2
Hits on station 2
10 modules, DSSD 63×63 mm2
Hits on station 4
18 modules, DSSD 63×63 mm2
#41
Hits on station 3
14 modules, DSSD 63×63 mm2
#13
Module | Chip
#41 | 5Y
#14 | 5X
#9 | 5X
#39 | 1X
#19 | 5X
#13 | 4X, 5X
#17 | 4X, 5X
#12 | 4X, 5X
#5-21 | 5X
#23 | 4Y
Run | 8305
Beam | Xe
Energy | 3.8 Gev/n
Target | Csl (2%)
D.D. Chemezov

## Slide 4: FEE board replacing
Before replacement
After replacement
To eliminate the dead zones it was necessary to assemble a new FEE board completely, 5 ASICs were spent to eliminate each of defects. The old boards were disconnected from the modules and replaced with a new ones. Total: 9 new black boards (p+ strips) were built, 45 VATAGP 7.2 ASICs were used (30 spare ASICs available).
Bad zone Chip #5
Mod #17
Mod #17
Yu.A. Kopylov; O.G. Tarasov

## Slide 5: Repairable FEE boards with new compound
Before replacement
Ultra Light-Weld 9008 Flexible, UV-Curable Encapsulant
Does not create an additional optical shield
Can be mechanically removed so FEE boards are repairable
Black BE-08 encapsulant (polarization temperature 100℃)
an additional optical shield for chips
Not removable mechanically or with chemical solvents (dimethyl sulfoxide, dimethylformamide, CH2O2)
Replacement of one chip:
removed chip
E.A. Streletskaia, Yu.A. Kopylov; O.G. Tarasov
After replacement

## Slide 6: Cosmic tests of FSD
DAQ
Caen (HV)
FUG (LV)
Trigger signal discriminator
Dimensions:
Scintillator: 150x600 mm2
Station-1, half-plane: 93x180 mm2
Station-2, half-plane: 126x300 mm2
Station-3, half-plane: 126x420 mm2
Station-4, half-plane: 126x540mm2
E.V. Martovitsky; S.V. Khabarov
Si-Plane
Cooling system
Scintillator

## Slide 7: FSD: after replacing
Cosmic tests station 1
93643 hit entries
111664 hit entries
Cosmic tests station 2
201548 hit entries
198035 hit entries
All dead zones have been eliminated!
124789 hit entries
138624 hit entries
Cosmic tests station 4
165513 hit entries
319725 hit entries
Cosmic tests station 3
E.V. Martovitsky; D.D. Chemezov
Beam hole

## Slide 8
#1-20
#37
#3/2
#5/2
#7/2
#9/2
#9/1
#7/1
#5/1
#3/1
#4-0/22
#6-0/22
#5-0/22
#1-0/22
#2-0/22
#21
#24
#41
#16
#34
#32
#27
#42
#38
#35
#3-0/22
#4-21
#6-21
#2-21
#22
#7
#29
#15
#30
#33
#9
#36
#14
#17
#13
#10-2017
#19
#10
#20
#28
#9-2017
#39
#11
#26
#23
#40
#31
#18
#25
#5-21
#12
beam
Station 1
Station 4
Station 3
Station 2
0
1
2
3
4
5
0
1
2
3
4
5
6
7
8
9
0
1
2
3
4
5
6
7
8
9
10
11
12
13
0
1
2
3
4
5
6
9
10
11
12
13
14
15
7
8
16
17
E.V. Zubarev
Repaired modules (10 pcs)

## Slide 9
*-position of the origin point of the Si-sensors in the coordinate plane 0XY (tied to the outer reference pins).
**-the distance from the base surface of the base bar to the middle of the Si-sensors surface farthest from it.
reference pin
reference pin
base bar
Si-detectors
Inclined strips side (N+)
Straight strip side (P+)
Coordinates of modules (Si-sensors) after disassembly-assembly FSD-planes.
Difference in Si-sensors position data in half-plane # 5/2 (25.03.24 compared to 06.07.22)
Position of Si-sensors in half-plane
# 5/2 (25.03.24) after repair
base surface
Sensor position | X±0.02* (mm) | Y±0.02* (mm) | Turn in 0XY plane (deg.)** | Z±0.2*** (mm) | Module serial number
Si-1a | 65.55 | 164.23 | -0.03 | 16.0 | #34
Si-1b | 65.59 | 101.25 | 0.01 | 15.7 | 
Si-2a | 125.45 | 164.21 | 0.04 | 27.6 | #16
Si-2b | 125.55 | 101.26 | -0.01 | 27.6 | 
Si-3a | 248.63 | 134.21 | -0.04 | 14.0 | #41 repair
Si-3b | 248.57 | 71.22 | -0.03 | 13.9 | 
Si-4a | 308.42 | 164.24 | -0.01 | 25.8 | #24
Si-4b | 308.46 | 101.26 | 0.0 | 25.9 | 
Si-5a | 368.64 | 164.25 | -0.06 | 14.2 | #21
Si-5b | 368.56 | 101.26 | -0.03 | 14.0 | 
Sensor position | ΔX (mm) | ΔY (mm) | Changin of turn (deg.)** | ΔZ (mm) | Module serial number
Si-1a | -0.08 | -0.02 | 0.0 | 0.1 | #34
Si-1b | -0.05 | -0.01 | 0.0 | 0.0 | 
Si-2a | -0.09 | -0.02 | 0.0 | 0.1 | #16
Si-2b | -0.06 | 0.0 | 0.0 | 0.1 | 
Si-3a | -0.06 | -0.02 | -0.02 | -0.5 | #41 repair
Si-3b | -0.03 | -0.01 | 0.0 | -0.3 | 
Si-4a | -0.08 | -0.01 | 0.0 | -0.3 | #24
Si-4b | -0.05 | 0.0 | 0.0 | -0.1 | 
Si-5a | -0.07 | -0.02 | 0.01 | -0.2 | #21
Si-5b | -0.05 | -0.01 | 0.0 | -0.2 | 
Changing of Si-sensors positions after disassembly and reassembly does not exceed:
- 100 μ‎m for X-coordinate
- 50 μ‎m for Y-coordinate
E.V. Zubarev

## Slide 10
FSD: New coordinates of Si-sensors in the plane
Position of Si-sensors in plane
# 5/2 (25.03.24) after repair
Position of Si-sensors in plane
# 7/1 (27.03.24) after repair
Position of Si-sensors in plane
# 9/2 (17.04.24) after repair
Position of Si-sensors in plane
# 9/1 (10.04.24) after repair
- New coordinates were entered into the data table for positions of detectors in FSD-planes
Measurement of Si-detectors position coordinates in 2024y were carried out on a new tooling for plane fixation on the table of video-measuring microscope "NORGAU" NVM II-5040D (3 points of support).
Sensor position | X±0.02* (mm) | Y±0.02* (mm) | Turn in 0XY plane (deg.)** | Z±0.2*** (mm) | Module serial number
Si-1a | 65.55 | 164.23 | -0.03 | 16.0 | #34
Si-1b | 65.59 | 101.25 | 0.01 | 15.7 | 
Si-2a | 125.45 | 164.21 | 0.04 | 27.6 | #16
Si-2b | 125.55 | 101.26 | -0.01 | 27.6 | 
Si-3a | 248.63 | 134.21 | -0.04 | 14.0 | #41 repair
Si-3b | 248.57 | 71.22 | -0.03 | 13.9 | 
Si-4a | 308.42 | 164.24 | -0.01 | 25.8 | #24
Si-4b | 308.46 | 101.26 | 0.0 | 25.9 | 
Si-5a | 368.64 | 164.25 | -0.06 | 14.2 | #21
Si-5b | 368.56 | 101.26 | -0.03 | 14.0 | 
Sensor position | X±0.02* (mm) | Y±0.02* (mm) | Turn in 0XY plane (deg.)** | Z±0.2*** (mm) | Module serial number
Si-1a | 65.47 | 164.14 | 0.03 | 28.6 | #14 repair
Si-1b | 65.48 | 101.11 | 0.04 | 28.0 | 
Si-2a | 125.39 | 164.27 | 0.06 | 16.2 | #36
Si-2b | 125.48 | 101.29 | 0.05 | 15.8 | 
Si-3a | 185.42 | 164.13 | 0.05 | 28.0 | #9 repair
Si-3b | 185.47 | 101.12 | 0.04 | 27.8 | 
Si-4a | 308.46 | 134.28 | -0.04 | 14.2 | #33
Si-4b | 308.41 | 71.31 | 0.03 | 14.0 | 
Si-5a | 368.43 | 164.21 | 0 | 25.8 | #30
Si-5b | 368.41 | 101.21 | 0.02 | 26.1 | 
Si-6a | 428.48 | 164.27 | -0.02 | 14.5 | #15
Si-6b | 428.46 | 101.27 | -0.01 | 14.2 | 
Si-7a | 488.57 | 164.18 | -0.06 | 25.7 | #29
Si-7b | 488.50 | 101.18 | -0.06 | 26.0 | 
Sensor position | X±0.03* (mm) | Y±0.05* (mm) | Turn in 0XY plane (deg.)** | Z±0.2*** (mm) | Module serial number
Si-1a | 65.58 | 164.27 | 0.01 | 16.2 | #39 repair
Si-1b | 65.59 | 101.28 | 0 | 15.8 | 
Si-2a | 125.45 | 164.23 | 0.05 | 27.2 | #9-2017
Si-2b | 125.49 | 101.11 | 0.04 | 27.4 | 
Si-3a | 185.41 | 164.22 | 0.07 | 16.2 | #28
Si-3b | 185.49 | 101.22 | 0.06 | 15.9 | 
Si-4a | 245.51 | 164.29 | 0.03 | 27.8 | #20
Si-4b | 245.53 | 101.28 | 0.01 | 27.7 | 
Si-5a | 368.44 | 134.32 | 0.01 | 14.1 | #10
Si-5b | 368.43 | 71.32 | 0.02 | 13.9 | 
Si-6a | 429.39 | 164.32 | 0.01 | 26.5 | #19 repair
Si-6b | 428.39 | 101.33 | 0.02 | 26.2 | 
Si-7a | 488.58 | 164.31 | -0.06 | 14.5 | #10-2017
Si-7b | 488.51 | 101.30 | -0.04 | 14.2 | 
Si-8a | 548.50 | 164.31 | -0.01 | 25.9 | #13 repair
Si-8b | 548.47 | 101.29 | -0.01 | 25.9 | 
Si-9a | 608.62 | 164.29 | -0.07 | 14.2 | #17 repair
Si-9b | 608.54 | 101.28 | -0.04 | 14.0 | 
Sensor position | X±0.03* (mm) | Y±0.05* (mm) | Turn in 0XY plane (deg.)** | Z±0.2*** (mm) | Module serial number
Si-1a | 65.46 | 164.28 | 0.03 | 15.0 | #12 repair
Si-1b | 65.59 | 101.21 | 0.05 | 15.3 | 
Si-2a | 125.48 | 164.31 | 0.04 | 26.4 | #5-21 repair
Si-2b | 125.56 | 101.18 | 0.02 | 27.2 | 
Si-3a | 185.54 | 164.36 | 0.01 | 15.2 | #25
Si-3b | 185.60 | 101.35 | 0.00 | 15.3 | 
Si-4a | 245.39 | 164.34 | 0.06 | 26.7 | #18
Si-4b | 245.50 | 101.13 | 0.06 | 27.2 | 
Si-5a | 368.46 | 134.37 | -0.02 | 14.4 | #31
Si-5b | 368.48 | 71.37 | 0.01 | 14.0 | 
Si-6a | 428.42 | 164.35 | 0.00 | 25.8 | #40
Si-6b | 428.48 | 101.35 | 0.01 | 25.9 | 
Si-7a | 488.61 | 164.35 | -0.03 | 13.7 | #23 repair
Si-7b | 488.59 | 101.34 | -0.05 | 13.7 | 
Si-8a | 548.45 | 164.37 | 0.01 | 25.7 | #26
Si-8b | 548.49 | 101.38 | 0.01 | 25.8 | 
Si-9a | 608.60 | 164.32 | -0.04 | 14.7 | #11
Si-9b | 608.60 | 101.29 | -0.02 | 14.3 | 
E.V. Zubarev

## Slide 11
Added a 3rd geodetic marker to control the inclination of the plane around the X axis
Screws to eliminate rotation of the plane around the Z-axis
(2022 Xe-run:
∠= 0,11° ÷ 0,22°)
Refinement of FSD-planes positioning device in the BM@N experemental hall.
Y
X
Y
Adjustment shims to eliminate plane rotation around the Y axis.
(2022 Xe-run:
∠= 0,11° ÷ 0,74°)
Z
X
Y
Z
X
Z
E.V. Zubarev

## Slide 12: Conclusion
Eliminated defects in FSD-planes, all defects were related to FEE-chips failures. At the end of Xe-run (February 2023) number of problematic chips were 2,7% (from 480 pC.), at the moment - 0%.
FSD-planes are ready to installation and alignment in BM@N experimental hall .
True coodrinates of Si-detectors in FSD half-planes after disassembly-assembly during defects repair were measured with accuracy ± 50μ‎m.
Refined FSD-planes positioning devices for alignment in BM@N experimental hall (added a 3-rd geodetic marker).

## Slide 13
Thank you for your attention!

## Slide 14
Additional slides

## Slide 15: SiBT station #3 status
49 and 50th strips are shorted out on the detector. Irreparable defect
#49
#50
There was no contact on the connector. The defect was eliminated by cleaning the connectors
 | Imax, nA | IΣ, μA
May 2023 | 182 | 7.1
April 2024 | 219 | 5.9
Delta | ≈40 | 1.2

## Slide 16: SiBT FEE during Xe-run
Chip 2 (64-127 ch.)
Amplitude distribution (SiBT #2 n+ side)
Channel distribution of amplitudes (SiBT #2 n+ side)
Hit plot SiBT #2
This slide shows SiBT FEE operation in XE-run on example of SiBT #2 (run 8387, 1.02.2023). The gains of FEE-chips on n+ side were not tuned for signals on the order of 10.9 pC (energy losses of Xe with energy 3.8 GeV/n are 245.5 MeV). Because of this signals from one of FEE-chips were below th threshold and were not written (in this case – chip 1 (channels 0 – 63)
pileups
Chip 1 (0-63 ch.)
Chip 1
Chip 2
Notes: На этом слайде показано почему в течение сеанса на бим трекерах «отваливались» чипы. Правая верхняя картинка – распределение сигналов с каждого из стрипов. Можно видеть, что первые 64 канала (первый чип) почти полностью обрезан порогом. Те полосы, которые остались на графике – временные наложения разных событий. Основная часть сигнала должна была бы быть в области ниже 300 отсчётов АЦП, но это значение «попало за порог».
Нижняя правая картинка – амплитудный спектр для того же трекера. Можно видеть в начале правый хвост Ландау-распределения сигналов с первого чипа (который обрезан), по центру – Ландау для второго чипа. Маленький пик справа обусловлен временными наложениями.
На левой нижней картинке показано распределение координат попадания ксенона в бим трекер. Половина «отвалилась» - загрузка значительно меньше, чем у другой половины

## Slide 17: SiBT FEE status
The gain spread of ASICs (input 10 pC, ts=240 ns+120 ns)
Linearity of ASICs after PCB tuning (ts=240 nS)

## Slide 18: FST Si module test results
3
2
RA source positions
1 2 3
17
2
1
3
n+ side p+ side
639
0
0
639
Measurement with RA source cross-section
(trigger rate: 150 Hz)
83
Electron end-point energy up to 3.27 MeV (214𝐵𝑖)
1
Module ID | <ENCp+> ē RMS | <ENCn+> ē RMS | SNR p+ side | SNR n+ side | Bad channels ratio, %
6_0 | 1963.60 | 2 237.88 | 12.22 | 10.72 | 1.25
HV = 75 V; T = 17 ⁰C
HV = 75 V; T = 17 ⁰C
HV = 75 V; T = 17 ⁰C
Ufd = 30 V

## Slide 19: Измерения положения Si-детекторов в модулях и плоскостях проводились на бесконтактном видеоизмерительном микроскопе «NORGAU» NVM II-5040D.
«NORGAU» NVM II-5040D | 
Диапазон перемещений по осям X и Y (мм) | 500 х 400
Диапазон перемещений по оси Z (мм) | 250
Допускаемая абсолютная погрешность линейных измерений по осям X и Y (мкм)* | ±(2.5+L/200)
Допускаемая абсолютная погрешность линейных измерений по оси Z (мкм)* | ±(2.5+L/100)
* L –измеряемая длина в мм | 
видеоизмерительный микроскоп
стол оператора
измеряемый объект
(плоскость из 5 модулей)
объектив
предметный стол
отсутствие сварки
Si #1
Si #2
реперный крест детектора
Топологические знаки на детекторе можно использовать в качестве реперов
измеряемый объект

## Slide 20: Beam profilometer is necessary for beam tuning (alignment of the center beam with the center of the target)
detector: DSSD, (128p+×128n+), strips pitch = 475 μm, thickness (Si) -175 μm, active area (61 × 61) mm2. Strips are combined in pairs on detector board. Total have (64p+×64n+), pitch = 950 μm
mechanical design: the plane of the profilometer is automatically removed from the beam zone to the parking position
FEE based on two VA32HDR11 ASICs. Total number of channels are 64 for X and 64 for Y coordinate. Dynamic range
-35pC - +25pC. Peaking time 800ns. Self trigger
mode based on TA32cg2 ASICs
current status:
FEE have been designed, manufactured and tested.
Next step is testing FEE with Si detector placed in
flanges with alpha-source (5.5 MeV) and
standalone DAQ subsystem.
Beam profilometer upgrade
working position
New readout electronics for beam profilometer
Beam tracker Cross-board
Beam profilometer
standalone DAQ