## Slide 1
Spectrometer Complex of the IBR-2 High Flux Pulsed Reactor: Current State and Recent Developments
D.P. Kozlenko
Frank Laboratory of Neutron Physics, Joint Institute for Nuclear Research, 141980 Dubna, Russia

## Slide 2
SPECTROMETER COMPLEX OF IBR-2
14 instruments for condensed matter research,
2 in development stage
Diffractometers:
HRFD, DN-6, RTD, DN-12, FSD, SKAT, Epsilon, FSS
Small Angle Scattering Spectrometer: YuMO
Reflectometers:
REMUR, REFLEX, GRAINS
Inelastic Neutron Scattering Spectrometer:
NERA
New Instruments in development:
Small Angle Neutron Scattering and Imaging (Radiography and Tomography) Spectrometer
BJN (INS in inverted geometry)
Radiography and Tomography Spectrometer:
NRT

## Slide 3
1. Condensed Matter Physics and Materials Science,
2. Physics of Nanosystems and Nanoscale Phenomena,
3. Physics of Complex Liquids and Polymers,
4. Biophysics and Pharmacology,
5. Applied Materials and Engineering Sciences.
Main research directions :

## Slide 4
High resolution Fourier Diffractometer
For precise studies of crystal structure and microstructure of materials
Type of detector Scattering angles range | scintillation ZnS(Ag) with time-focusing geometry 2θ = (133−175)°.
Total solid angle | ∼2.0 sr (12.5 times greater than the existing one)
Basic Parameters
Development of the backscattering detector system with a large solid angle
Sub-project 04-4-1149-3-1-2021/2028 “Construction of a wide-aperture backscattering detector (BSD-A) for the HRFD diffractometer”, leader V.M.Milkov

## Slide 5
Fourier Stress Diffractometer
for studies of internal stresses in constructional materials and industrial products
Beam profiles:
analytical calculations and Monte Carlo simulation
Gauge volume of 1.8 mm
New Fourier Chopper (Airbus Defense and Space), fabricated in the end of 2020)
High-resolution RTOF diffraction pattern measured with a new Fourier chopper on a standard α-Fe sample
The expected increase in the intensity of registered neutrons is about 1.5 times.
ASTRA detector system: 14 new modules were installed at scattering angles 2θ = ±90°
Arrangement of sensitive ZnS(Ag) layers of the ASTRA 90°-detector in the horizontal scattering plane
3D model of 90° ASTRA detector made of 7 elements with time focusing geometry)

## Slide 6
DN-6 Diffractometer
for studies of crystal and magnetic structure of materials at ultrahigh pressures up to 40 GPa
Insertion of a new curved neutron guide
section with m=3 side coating
into the beam delivery system
of DN-6 diffractometer
Replacement of the central part of the neutron guide (from m = 1 to m = 3, made by SwissNeutronics)
Diamond anvil cells (DAC), Almax plate (left) and home-made (right) design
Calculated dependence of the gain factor
of the flux density on the neutron wavelength
for different coating for side mirrors
of curved neutron guide: m2, m3 vs m1
Neutron diffraction patterns of LaMnO3 measured at pressures up to 39 GPa
Minimal sample volume: V ~ 0.02 - 0.05 mm3
Upgrade of the detector system: new forward scattering bank (2θ=32-54°)
General view of the upgraded detector system, the new detector bank on the right . It contains 15 independent modules
arranged in a ring. Design of one module composing of 6 independent 3He neutron counters
The upgraded detector system will provide new opportunities for studies of magnetic structures of materials in the extended d-spacing range 0.8 – 12 Å.
Expected intensity gain  2 times

## Slide 7
DN-12 Spectrometer
for studies of crystal and magnetic structure of materials at high pressures up to 7 GPa and microsamples
DN-12 spectrometer (view before Dec 2024)
Sapphire anvil cell
Upgrade of the detector system and sample environment:
Installation of the multisectional 3He detector module (2θ=90°), consisting of 96 independent counting elements in the common gas volume.
Upgrade of the forward scattering detector module (2θ=45°) by installation of 2 additional circular arrays of individual 3He counters (16 counters in each array) to available one.
Development of new cryostat based on CCR (min T ~ 5 K)
The expected increase in the intensity of registered neutrons is up to 6 times (2θ=90°) and 2-3 times (2θ=45°).
The expected increase in the accessible pressure range from 7 to 10-12 GPa.
The expected reduction of the minimal sample volume from 1.0 to 0.2-0.5 mm3.
Detector system of DN-12
Detector modules at 2θ=90 (left) and 45° (right)
New cryostat

## Slide 8
Operated before 2022 in collaboration with the Bonn University, Prof. N.Froitzheim group, Germany)
Crystallographic Preferred Orientations (CPO) in eclogites
SKAT Diffractometer
for texture analysis of earth rocks, materials, biological and paleontological objects
Typical sample measurement time: 24-48 h.
Measurement demands exceed possibility in 2-3 times.
Installation of an additional detector module behind the existing one with the same scattering angle and geometry and split of neutron beam into two parts.
The electronics for the second module is fabricated, the manufacturing of detector pre-amplifiers and installation of neutron counters is planned in 2025.
The second goniometer unit is installed.
Two samples could be measured simultaneously in two experimental positions.
The number of experiments will be increased in 2 times.

## Slide 9
Upgrade for texture analysis:
With minor changes of orientation of the detector system elements with respect to the sample position Epsilon could be used for texture measurements with a pole figure measurement grid varied from 15 x 15 to 7.5 x 7.5 deg.
Replacement of old 3He counters (being in operation for about 25 years) by new ones. Expected intensity gain up to 2 times.
EPSILON Diffractometer
Previous use: stress and strain measurements in earth rocks and construction materials.
Demand for experimental time was insufficient.
Most Users were from Germany and Poland.
Extension of research program: upgrade for texture measurements (geophysics, cultural heritage, materials science).
Operated before 2022 in collaboration with the KIT, Prof. F.Schilling group, Germany)
Epsilon spectrometer (view before Nov 2024)
Calculated positions of detector modules for texture measurements

## Slide 10
Small-Angle Neutron Scattering and Imaging Spectrometer
Beam size | 5050 mm2
Wavelength range: | 0.5 – 10 Å
q-range: | 0.003 – 0.5 Å-1
Angular resolution | 5 - 20 %
Max. sample size | 5050 mm2
Flux at sample | 2.0106 cm-2 s-1
Detector, PSD | 3He, 7070 cm2, resolut. 55 mm2
Expected parameters
Current State
Installed at IBR-2
Vacuum tank for PSD
L = 2 m, R = 14 .3 m, N = 20 ( m = 2)
Delivered to JINR, in preparation for installation
Delivered to JINR
Changeable beam forming system
D = 1.6 m, L = 12 m
L = 6 m
Bender
PSD
Delivery is delayed, considered for 2025
Produced by Mirrotron (Hungary)

## Slide 11
Progress in Development of New Inelastic Scattering Spectrometer in Inverse Geometry BJN
Schematic layout of BJN
Bell-shaped analyzers composed of sets of highly oriented pyrolytic graphite (HOPG) single crystals
Detector system: two assemblies of cylindrical 3He counters d=18 mm (analogue of SNM-18), 5 counters each.
A prototype element of the focusing analyzer of the new inelastic neutron scattering spectrometer BJN was designed and fabricated, delivered to Dubna (Frako-Term, Poland).
The 3He counters for the detector system are purchased and development of the detector system is in progress.
250 HOPG crystals were purchased and their quality was tested at IR-8 reactor (NRC KI).
The prototype element of the focusing analyzer will be equipped with the HOPG crystals and test measurements will be planned at one of the available neutron beams of IBR-2.

## Slide 12
Thank You for Your Attention!