## Slide 1
Development of a temperature controller for cryogenic systems at the NICA accelerator complex
Dzugaev Maxim
Technic, LHEP JINR
«ALUSHTA 2025»

## Slide 2
To conduct experiments on the BM@N installation, targets with a unique ratio of the working substance and the substance from which the target is made were needed. The requirement to have a minimum number of background events is a basic requirement for precision research. To implement the set physical program, a new cryogenic target is being created with hydrogen, deuterium and helium in a liquid state using cryocooler cooling. One of the main tasks in creating a cryogenic target using cryocooler cooling is to maintain a set temperature in the mode of recondensation of evaporating vapors of liquid hydrogen, deuterium or helium and a constant level of these liquids in the vessel.
A long time ago…
in 2022
t = 20,28 K

## Slide 3
1-st generation of temperature controller
The refrigeration unit in use does not have a power control system and operates continuously. In the mode of hydrogen liquefaction and target filling, the entire refrigeration capacity produced by the cryocooler is used. Further, in the mode of maintaining the level of liquid hydrogen, much less refrigeration power is required, Excessive cooling capacity can cause hydrogen crystallization. The task of maintaining the set temperature and liquid level is solved by compensating for the excessively produced refrigerating power by an electric heater.

## Slide 4: The main block diagram
3U 19 inches
Power supply 500W
Display
RS-485
Meas Module (2 chanells)
~220V
Heater Module
(< 100W)
Master Module
Load
Ethernet
Temperature sensors

## Slide 5
Fan control
Wirenboard
Display
Fans
RS-485
Power Supply 60V -> 12V & 5V
Ethernet
Master Module
Heater
Meas
Master

## Slide 6: Heater Module
p
MCU
ADC
Power Supply 60V -> 12V & 5V
Pulse converter
PWM
Linear converter
DAC 0-5V
Current measure
< 48V
< 45V
Load output
< 100W
    Temperature stabilization accuracy – 1 mK ± 0,1 %.
    Channel power – to 100 W.
    Maximum output voltage – 50 V.
    Maximum output current – 2 А.
    Accuracy of measurement of output current, voltage – not worse ± 0,3 %.
    Recommended heater resistance – from 25 to 100 Ohm.

## Slide 7: Meas Module
p
MCU
DAC 0-5V
DAC 0-5V
Current source 100n-1mA
Current source 100n-1mA
ADC
I+
I-
V+ V-
R1
I+
V+ V-
R2
I-
The range of measured temperatures is from 1.5 to 310 K (determined by the type of sensor).
The relative error of temperature measurements does not exceed ± 0.1% (determined by the sensor error in a certain range).
The bit depth of the ADC is 24 bits.
The measurement frequency is 10 Hz.
The range of the current source is from 100 nA to 1 mA.
The relative error of the current source does not exceed ± 0.1% in the range to 1 µA.

## Slide 8: Why module system?
The modular system allows you to assemble different configurations for different needs (tasks) In the standard version, the 3U 19 inches crate can accommodate 3 meters and 3 heaters, the number of which may vary.

## Slide 9: It’s not a model, it’s real !

## Slide 10
And we test it
PT-100 |  |  |  | 
Resistance |  | Temperature |  | Delta
Lakeshore | ККВМ | Lakeshore | ККВМ | 
3,4100 | 3,4107 | 28,036 | 28,026 | 0,01
7,9100 | 7.90675 | 46,006 | 45,99 | 0,016
32,0800 | 32,0800 | 105,02 | 104,998 | 0,022
63,0100 | 62.98445 | 180,09 | 180,011 | 0,079
131,6100 | 131.55432 | 354,99 | 354,834 | 0,156
182,0300 | 182.04251 | 489,97 | 490,02 | -0,05
247,7000 | 247.69421 | 674,98 | 674,958 | 0,022
293,0800 | 293.08706 | 809,98 | 809,79 | 0,19
PT-1000 |  |  |  | 
Resistance |  | Temperature |  | Delta
Lakeshore | ККВМ | Lakeshore | ККВМ | 
1020,0000 | 1019.5772 | 278,26 | 278,164 | 0,096
1099,5000 | 1099.5483 | 298,83 | 298,687 | 0,143
235,2000 | 235.19552 | 85,009 | 84,987 | 0,022
99,2000 | 99.27473 | 51,992 | 51,953 | 0,039
51,4000 | 51.41354 | 35,979 | 35,973 | 0,006
34,1000 | 34.09416 | 28,027 | 28,021 | 0,006
2477,1000 | 2477.11177 | 675,01 | 675,003 | 0,007

## Slide 11
Some graphic

## Slide 12
More graphics

## Slide 13
Conclusion

## Slide 14
Thank you for your attention!