## Slide 1: Waiting for the beam of relativistic nuclei 124Xe at BM@N N. Zamyatin on behalf of Forward Silicon Detector team
14th Collaboration Meeting of the BM@N Experiment at NICA, 13-15 May 2025

## Slide 2: 1. FSD на канале BM@N для сеанса май-июнь-2025
station 1
station 2
201548 hit entries
198035 hit entries
124789 hit entries
138624 hit entries
station 4
165513 hit entries
319725 hit entries
station 3
93643 hit entries
111664 hit entries

## Slide 3: FSD Si module test results (module #41)
β-source positions
1 2 3
17
n+ side
p+ side
639
0
0
639
2
3
Sigma values of each FSD Si module channels at 75 V
1
2
3
#1 #2 #3 #4 #5
Occupancy distributions in FSD Si module channels after noise suppresion
1
Module view after US- Bonding detectors and PAs, number of US- Bonding – 640, Ø Al wire – 25 μm
Module ID | Dark current (50V), nA | Mean noise (+p), ch.ADC | Mean noise (+n), ch.ADC | MPV (+p), ch.ADC | MPV (+n), ch.ADC | S/N p+ side | S/N n+ side | Bad channels ratio, %
41_0 | 3 584,00 | 36,93 | 50,76 | 536,20 | 578,89 | 14,09 | 13,32 | 0,55
1ch. ADCp+ = 45 e, 1 ch.ADCn+= 42 e

## Slide 4
2. Пучковый трекер SiBT1÷SiBT3 к сеансу 2025
Радиационный автограф
пучка 124Xe(3.8ГэВ/н) за
сеанс 2022-2023г.

## Slide 5
Прогноз на ресурс у ДССД детекторов при интеграле ядер за сеанс N = 4.44×1010:
124Xe(3,8 ГэВ/н): 10 сеансов (NXe = 4.44×1011), Istrip(max) = 2 мкА;
197Au(3,8 ГэВ/н): 5 сеансов, Istrip(max) = 2 мкА;
209Bi(3,8 ГэВ/н): 4 сеанса, Istrip(max) = 2 мкА .

## Slide 6
5mm
5mm
target
FSD
GEM
Si- pin-detector (rad_monitor) :
- Detectors size: (5×5×0,3) mm3
- Detectors thickness: 300 μm
- Active area: 13 mm2
ΔI=αI×V×Ф
ΔI = Iф − I0, (A):
Iф – dark current after irradiation;
I0 - dark current before irradiation;
αI = (5±0,5)·10−17 (A×cm-1) – radiation damage constant-Si (for neutrons 1MeV at +20°);
V, (cm³) – volume of SCR (spice charge region at Ufd);
Ф, (cm-2) – eq. neutron fluence (1 MeV)
Z
Y
Si-det.
FEE
3. Radiation Si-monitors on FSD region for equivalent neutron fluence measurement

## Slide 7
target
Z
Station #4
Station #1
beam
Y
X
Si-monitors
4
5
6
3
1
2
# Si-monitors | L -расстояние от центра мишени до центра Si монитора (мм)
1 | 156
2 | 137
3 | 156
4 | 197
5 | 190
6 | 197
7 | 415
8 | 432
9 | 415
10 | 532
11 | 438
12 | 532
7
10
8
9
11
12
Radiation Si-monitors on FSD region for equivalent neutron fluence measurement
L

## Slide 8
beam
Si-монитор:
Sactiv – 13mm2
4. Облучение LYSO (ФИАН) crystals
Forward hadron calorimeter FHCal
Beam
LYSO crystals 8х10х4мм3
Кристаллы Lu2SiO5-z:Y3+:Ce3+:Ca2+
(Lu1.961Y0.037Ce0.00068Ca0.00038Si1O5-z, z=0.00038.)
По инициативе коллег из ФИАН:
Козлов В.А., Завертяев М.В.
и согласованию с сотрудниками ЛФВЭ:
Капишин М.Н., Замятин Н.И. было принято решение облучения на пучке ядер Хе установки BM@N сцинтилляционных кристаллов LYSO

## Slide 9
Run 444 (10.02.2025)
Run 542 (28.04.2025)
5. Modernization of power grids on BM@N experimental hall (bld.205)
(March 2025)
Before modernization of power grids:
tmeasure = 22°
After modernization of power grids:
= 1949
1589
tmeasure = 26°

## Slide 10: Conclusion
FSD is ready for data taking with relativistic 124Xe beam
Coordinate planes position accuracy is 0.3 mm with XYZ axes
SiBT#1÷ SiBT#3 are ready for operation in the direct beam of heavy ions
Level of the FEE FSD pedestal noises did not increase after modernization of power grids (for Semen Piyadin)
Applied R&D:
measurements of the radiation background on the area of FEE and Si-detectors of FSD
Irradiation with Xe ions of the LYSO-crystals (ФИАН)