## Slide 1
Status of the Si-detectors on BM@N
N. Zamyatin on behalf of Si-detectors team, 14÷16 October 2025, 15th – CM/BM@N, Dubna
- Beam Profilometer;
- Beam Tracker (BT1 ÷ BT3);
- Forward Si-Tracker (FSD)
*Отчет 26.01.2024г
« Измерение фактического положения детекторов установки BM@N и их юстировка в проектное положение относительно пучка.»
ООО «Промышленные измерения»

## Slide 2: Beam profilometer is necessary for beam tuning in self trigger mode (alignment of the center beam with the center of the target)
detector: DSSD, (128p+×128n+), strips pitch = 475 μm, thickness (Si) -175 μm, active area (61 × 61) mm2. Strips are combined in pairs on detector board. Total have (64p+×64n+), pitch = 950 μm
mechanical design: the plane of the profilometer is automatically removed from the beam zone to the parking position
FEE based on two VA32HDR11 ASICs. Total number of channels are 64 for X and 64 for Y coordinate. Dynamic range of signals: -35pC ÷ +25pC. Charge of ionization for heavy ions Xe(10.8 pC/175 μm-Si), Au (20.9 pC/175 μm-Si), Bi) corresponded of ∆E-signals (240÷500 MeV/175 μm-Si).
Peaking time 800 ns.
Self trigger mode based on TA32cg2 ASICs
current status:
FEE have been designed, manufactured and tested with Si detector placed in flanges with alpha-source (226Ra) and standalone DAQ subsystem, Need to install and to align on beam pipe
1. New beam profilometer for heavy ions (Xe, Au, Bi)
working position
Beam profilometer inside beam vacuum station

## Slide 3: - a new coordinate plane (Fig.2) has been developed and manufactured for run 2025: DSSD with (128x128) strips, 175 µm thickness has been converted into a (64x64) strips configuration (combining the strips in pairs, for a total of 128 channels); - a new FEE variant based on the VA32HDR11 + TA32cg2 ASICs (there are ASICs available) has been developed, manufactured and tested with an alpha source, a large signal range of +/-20 pC/dE≈500 MeV (for more detail see report by D.Chemezov)
- Fig.1 shows the results of the 2022 tests Si profilometer version–1 with an alpha (4.5-7.6) MeV source; - the same results were expected when working with a 124Xe beam (dE/175 µm ≈ 245 MeV), but failed due to signal overlap (self trigger mode) and exceeding the dynamic range (ASIC-VA163/750 fC), both profilometers were removed to the "parking" position and were not used in the 2022-2023 run, in the future they can be used for light ions up to Ar;
Status of the new beam profilometer (64×64) strips for heavy ions (Xe, Au, Bi)
226Ra in three different positions
Fig.1.
Fig.2.

## Slide 4
2. Beam Tracker (BT1÷BT3) for Run-2025
Radiation autograph of the 124Xe beam(3.8 AGeV) from the 2022-2023 run.

## Slide 5
Prediction of radiation resource for Si detectors of SiBT with N = 4.44×1010 ions per run:
124Xe(3,8 AGeV): 10 runs (NXe = 4.44×1011), Istrip(max) = 2 µА;
197Au(3,8 AGeV): 5 runs, Istrip(max) = 2 µА;
209Bi(3,8 AGeV): 4 runs, Istrip(max) = 2 µА .
D.Chemezov

## Slide 6: 3. Forward Si-Tracker (FSD) for Run-2025
station 1
station 2
201548 hit entries
198035 hit entries
124789 hit entries
138624 hit entries
station 4
165513 hit entries
319725 hit entries
station 3
93643 hit entries
111664 hit entries
Mean noise of FSD half-stations from April 2025 up to November 2025

## Slide 7
Three type of Si-modules for Forward Si detector
63
93
FEE board
63
93
186
FEE board
63
63
126
FEE board
Detectors size: 63x63x0,3 mm3 (on 4” – FZ-Si-n wafers)
Topology: double sided microstrip (DSSD)
(DC coupling)
Pitch p+ strips: 95 μm;
Pitch n+ strips 103 μm;
Stereo angle between p+/n+ strips: 2.50
Number of strips/DSSD:
640 (p+)×614(n+)
Number of strips/module:
640 (p+)×640(n+)
Detectors size: 63×93×0,3 mm3 (on 6” – FZ-Si-n wafers)
Topology: double sided microstrip (DSSD)
(DC coupling)
Pitch p+ strips: 95 μm;
Pitch n+ strips 107,1 μm;
Stereo angle between p+/n+ strips: 2.50
Number of strips/DSSD:
640 (p+)×603(n+)
Number of strips/module:
640 (p+)×603(n+)
Station 2,3,4
DSSD 63x63mm2
42 modules
Station1
DSSD 63x93mm2
6 modules
New Station 5*
DSSD 63x93mm2
22 modules
Detectors size: 63x93x0,3 mm3 (on 6” – FZ-Si-n wafers)
Topology: double sided microstrip (DSSD)
(DC coupling)
Pitch p+ strips: 95 μm;
Pitch n+ strips: 107,1 μm;
Stereo angle between p+/n+ strips: 2.50
Number of strips/DSSD: 640 (p+)×603(n+)
Number of strips/module: 640 (p+)×640(n+)
For more details see report by Yu. Kopylov

## Slide 8: Plans and new developments on 2025-2027 years:
1. To develop and create 5th FSD station:
* DSSD/(93×63) mm2 – available;
* Pitch-adapters/640ch – available;
* ASIC-128ch/ENC0=150 rmse (JINR-Minsk) - manufacturing
stage of the 1st version;
* Station design – development stage;
* Si-module design – done;
* Pilot module – designed, manufactured, testing stage;
* FEE-PCB-640ch/new chip –development stage

## Slide 9: Half-plane of the new FSD-station #5 on the BM@N (11 modules with two Si-sensors (93х63)mm2 each)
Light-tight cooling air outlet
Cooling air inlet. t=+10°C
Tooling Hole for measuring the position of half-plane Si sensors
Tooling Hole for measuring the position of half-plane Si sensors
FEE cross-board
Si-detectors 93х63 mm

## Slide 10: Position of the new FSD-station #5 on the BM@N
Current position of GEM stations
Shifted position of GEM stations (Z+93 мм)

## Slide 11: Back up slides

## Slide 12: X-Z cross section of central tracker
Current position of GEM stations
Position of GEMs after shift (Z+93 мм)

## Slide 13
Текущее положение
станций GEM
Смещенное положение
станций GEM (Z+93 мм)