## Slide 1: The beam position measurement system of Booster NICA
E. Gorbachev, D. Monakhov, H.Nazlev,
S. Romanov, V. Volkov

## Slide 2: BPM systems
based on Beam Position Monitors (BPM), which are beam detectors located along the accelerator:
  Beam Position Monitor - Beam pickup with read-out electronics.
  4 or more BPMs per betatron oscillation period, typically located near each quadruple magnet.
deliver beam orbit information:
  Non-destructive monitoring based on the EM-field of the passing beam.
  turn-by-turn, bunch-by-bunch, or averaged over many turns
are a powerful beam diagnostics tool
  Machine commissioning, characterization of the beam optics, measurement of beam parameters, trouble-shooting

## Slide 3: Booster diagnostics layout
D
F
D
F
D
F
D
F
D
F
D
M
M
M
M
M
M
M
M
M
M
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpmZ
bpmX
bpm
bpm
bpm
bpm
bpm
Vacuum station
F quadrupole
D quadrupole
Magnet
Multipole corrector
BPM
scraper
One quadrant:
Lenses block:
BPM in lenses modules (X/Z) | 24
BPM in magnet modules (X/Z) | 20
BPM in bump region (X/Z) | 4
F
Total number of diagnostic devices in the ring:
PU(Z)
PU(X)
D
F

## Slide 4: Linear-cut pickup
X=a =
Pros:
Linear position response: easy displacement calculations.
High Transfer at low frequencies: high signal levels, good for long bunches, allows operations at low frequencies with high load impedance
Cons:
Complicated mechanics.
Limited to low frequency operation, typically <300 MHz.
High capacitive source impedance, reflective
*BPM Systems, Manfred Wendt, CERN,BI CAS 2018

## Slide 5: Booster pickup
Elliptical linear-cut electrostatic (“shoe-box”) pickup.

## Slide 6: Pickup linearity
Chart: 
Sum | 700 | 688 | 614 | 582 | 562 | 552 | 540 | 532 | 528 | 528 | 524 | 524 | 522 | 522 | 520 | 522 | 522 | 522 | 524 | 524 | 526 | 528 | 532 | 534 | 538 | 544 | 548 | 560 | 572 | 594 | 630 | 700 | 928
Delta | 344 | 332 | 278 | 242 | 214 | 192 | 168 | 148 | 128 | 108 | 92 | 72 | 54 | 38 | 20 | 6 | -14 | -34 | -52 | -68 | -86 | -104 | -120 | -138 | -158 | -176 | -196 | -216 | -240 | -266 | -302 | -360 | -504
Chart: 
Normalized | 0.49143 | 0.48256 | 0.45277 | 0.41581 | 0.38078 | 0.34783 | 0.31111 | 0.2782 | 0.24242 | 0.20455 | 0.17557 | 0.1374 | 0.10345 | 0.072797 | 0.038462 | 0.011494 | -0.02682 | -0.065134 | -0.099237 | -0.12977 | -0.1635 | -0.19697 | -0.22556 | -0.25843 | -0.29368 | -0.32353 | -0.35766 | -0.38571 | -0.41958 | -0.44781 | -0.47937 | -0.51429 | -0.5431
BPM prototype was measured with current pulses through movable wire and rods of different diameters.
Measurements show good linearity of the BPM prototype despite mechanical problems.
wire position
wire position
V, mV
∆/Σ

## Slide 7: Read-out electronics
The read-out system interfaces the BPM pickup to the accelerator data acquisition (control) system:
Signal conditioning,
Normalization,
Linearization,
Conversion to a digital format to provide a time-stamped beam position.

## Slide 8: Libera Hadron
Libera Hadron (Instrumentation Technologies, Slovenia):
Used in hadron synchrotrons
Bunch-by-bunch position measurements
Large buffers for ADC and position data storage
Tune measurements, FFT processing, closed orbit feedback functionality
Performance:
Typical Bunch-by-bunch RMS uncertainty, Uin = 1.5 V peak, pulse FWHM = 100 ns ~ 4 µm
8-hour stability (ambient temp. = T±1° C): < 2 µm
Max temperature drift (ambient temp. range: 20 to 30°C): ~ 2 µm/°C

## Slide 9: Libera Hadron block diagram
4 types of modules:
The beam position processor (BPM) module receives the signal from the pick-ups, does analog signal processing and conversion to digital. Contains the FPGA for digital data processing and RAM to store the data.
The inter connection board (ICB) provides both HW and SW interface towards the user’s control system. It performs the control actions of all boards in the Libera Hadron chassis and enables the data transfer to the user.
Timing (EVRX) module is reception of the events (including optical events) and the distribution of clock and trigger signals to all application boards.
Gigabit data exchange (GDX) module offers fast links (SFP connectors) through which the FA data are transferred in hard real time. These data can be used for building a fast feedback application.

## Slide 10: Libera Hadron Specifications
Specifications:
16 input channels (4 per module)
16 bit ADC
250MS/s
55 MHz input frequency range
6.5Gbps Rocket I/O
4GB memory per module
Data paths:
Broadband data (ADC@250MHz) -270MS (>1s of data), 4 channels
Bunch-by-bunch data – 270MS (>66s of data @1MHz bunch rep rate)
Synthetic beam position data – user defined decimation with averaging
Slow data stream (10 S/s)
Fast data stream (10kS/s)
Output data:
Raw(A,B,C,D)
Calculated beam position (X,Y)
Charge(SUM)
FFT, FFT peak

## Slide 11: BPM signals processing
The external preamplifier mounted close to the BPM is in charged to provide high intensity and low noise signals to the Libera Hadron unit. A huge 110 dB beam dynamic can be coupled with the Amplifier 110.
Supported gain ranges: 60, 50, 40, 30, 20, 10, 0, -10, -20, -30, -40 and -50 dB
Gain error between channels: ±0.1 dB
Gain switching time: <1 us

## Slide 12: BPM signals processing
Bunch-by-bunch algorithm, detection for “shoe-box” sensor:
Bunch detection.
Processing window calculation.
Baseline restorer (shoe-box).
Energy calculation.
Position calculation.
Output data rate = Bunch-by-bunch (varying)

## Slide 13: BPM signals processing
Timed BBB algorithm:
Bunch position synchronized with external RF clock.
RF clock can be delayed
Signal energy calculated in the gated region.

## Slide 14: BPM signals processing
The narrow-band analysis is a superheterodyne SDR receiver. The algorithm is intended for processing extremely low signals where individual bunches can not be detected. Four ADC signals are mixed to fIF with the DCO locked to external #RF signal using ADPLL :
SDR receiver.
Three filters enabling different output data bandwidths can be selected in the DSP nba block: iir0 – 1 kHz, iir1 – 2 kHz, iir2 – 5 kHz.
Output data rate = 50 kHz (fixed).
Narrow-Band Analysis “nba”

## Slide 15: #Start, #stop and #postmortem
The EVRX module supports both electrical and optical event based synchronization.
The #START and #STOP events relate to the instrument observation interval. The Observation cycles can be selected as being one-shot or repeatable. The #PM event can occur anytime. The Observation intervals are:
Raw ADC data storing: 2 GB ADC buffer
Bunch-by-bunch data storing: 2 GB Bunch-by-bunch buffer
Following events affect observation cycle:
#START (T2 trigger)
#STOP (T2 trigger)
#PM (T1 trigger)

## Slide 16
BPM signals processing
Different impedance matching schemes:
no matching (signal distorted).
buffer amplifier (S/N is affected).
matching transformer (signal amplitude is affected).
Matching with 6:1 transformer

## Slide 17: Booster beam parameters
Particles | Au31+
Nominal intensity | 2·109
Injection Energy | 3.2 MeV/n
Extraction Energy | 600 MeV/n
Harmonics number | 5 (after injection) / 1 (after cooling)

## Slide 18: RF parameters (1 stage)
 | I.1 | I.2 | I.3 | I.4 | I.5
Time, s | 0,0/0,05 | 0,05/0.1 | 0,1/0,4 | 0,4/0,45 | 0,45/0,475
B, T | 0,117/0,117 | 0,117/0,147 | 0,147/0,507 | 0,507/0,537 | 0,537/0,537
RF frequency, kHz | 587/587 | 587/736 | 736/2401 | 2401/2526 | 2526/2526
Bunches number | 0/5 | 5 | 5 | 5 | 5/0
Bunch RMS, m | -/5,3 | 5,3/3,75 | 3,75/3,64 | 3,64/5,6 | 5,6/-
Sync. Phase, deg. | 0/0 | 0/21 | 21/22,6 | 22,6/0 | 0/0
Peak current, mA | 1,16/3,68 | 3,68/6,54 | 6,54/21,9 | 21,9/15,0 | 15,0/5,0

## Slide 19: RF parameters (2 stage)
 | II.1 | II.2 | II.3 | II.4
Time, s | 0,0/0,05 | 0,05/0,1 | 0,1/1,1025 | 1,1025/1,1525
B, T | 0,537/0,537 | 0,537/0,567 | 0,567/1,77 | 1,77/1,80
RF frequency, kHz | 505,2/505,2 | 505,2/529,5 | 529,5/1110,2 | 1110,2/1117,4
Bunches number | 1/1 | 1/1 | 1/1 | 1/1
Bunch RMS, m | -/20,6 | 20,6/13,78 | 13,78/11,16 | 11,16/20,94
Sync. Phase, deg. | 0,0/0,0 | 0/31,2 | 31,2/42,1 | 42,1/0
Peak current, mA | 5,0/20,5 | 20,5/32,1 | 32,1/83 | 83/44,6

## Slide 20: BPM acquisition parameters
 | Stage I | Stage II
Peak current, mA | (1,16) 3,68 ÷ 21,9 | 5 ÷ 83
RF frequency, kHz | 587 ÷ 2526 | 505,2 ÷ 1117,4
Revolution period, us | 8.5 ÷ 2 | 2 ÷ 0,895
Bunch full length (±3σ), us | 1,28 ÷ 0,215 | 1,1 ÷ 0,285
Synchronous phase, deg. | 0 ÷ 22,6 | 0 ÷ 42,1
Number of bunches | 5 | 1

## Slide 21: Test bench
Pickup prototype with positioning system.
Libera Hadron.
Amplifier 110.
Arbitrary signal generator feeding a movable wire loaded by characteristic impedance.
Oscilloscope.
Main goal is to measure position and estimate the accuracy by using 3 algorithms with 2 different impedance matching schemes:
At nominal intensity injection (low current, low frequency)
At nominal intensity extraction (high current, high frequency)
At low intensity for beam commissioning.

## Slide 22: Impedance matching
Pickup on 50 Ohm load
Pickup loaded by
6:1 transformer
Pickup with preamplifier

## Slide 23: Libera hadron calibration methodics
Place wire to the geometric center of the pickup.
Libera ADC calibration: remove DC offset – automatic or manual.
Amplifier110 calibration: achieve equal gain of 4 channels in 12 gain ranges by achieving 0 position measurements in both planes.
  Take into account preamps or transformers gains.
  Obtain an even gain step over all ranges.
Calibrate SUM coefficient to obtain correct intensity measurements. The coefficient depends on gain.
Displace wire to the extreme position in one plane.
Calibrate Libera geometric coefficients to obtain correct position measurements in the plane.
Repeat 5,6 for the other plane.

## Slide 24: Position measurements results
 | BBB | BBB synth | tBBB | tBBB synth | NBA | NBA synth
Mean, mm | 4.857 | 4.853 | 5.210 | 5.184 | 4.986 | 5.028
Std err, um | 108 | 132 | 124 | 12.7 | 7.50 | 3.69
Preamplifiers, injection, 2e9 ions (3.7mA, 587kHz RF)
 | BBB | BBB synth | tBBB | tBBB synth | NBA | NBA synth
Mean, mm | 5.375 | 5.365 | 5.386 | 5.385 | 5.418 | 5.416
Std err, um | 547 | 221 | 131 | 13.6 | 7.63 | 4.01
Transformers, injection, 2e9 ions (3.7mA, 587kHz RF)
ADC with preamplifiers,
Amplifier110 gain=10dB
ADC with impedance transformers,
Amplifier110 gain=60dB

## Slide 25: Position measurements results
 | BBB | BBB synth | tBBB | tBBB synth | NBA | NBA synth
Mean, mm | 4.777 | 4.748 | 4.650 | 4.639 | 5.025 | 5.023
Std err, um | 1093 | 124 | 1200 | 113.1 | 74.2 | 38.4
Preamplifiers, injection, 2e8 ions (0.37mA, 587kHz RF)
 | BBB | BBB synth | tBBB | tBBB synth | NBA | NBA synth
Mean, mm | 1.646 | 1.761 | 3.925 | 3.891 | 5.392 | 5.404
Std err, um | 6382 | 650 | 776 | 96.8 | 88.3 | 37.3
Transformers, injection, 2e8 ions (0.37mA, 587kHz RF)
ADC with preamplifier,
Amplifier110 gain=40dB
ADC with impedance transformer,
Amplifier110 gain=60dB

## Slide 26: Position measurements results
 | BBB | BBB synth | tBBB | tBBB synth | NBA | NBA synth
Mean, mm | 4.976 | 4.975 | 5.060 | 5.051 | 4.958 | 4.957
Std err, um | 33.21 | 3.83 | 35.70 | 4.43 | 1.911 | 1.06
Transformers, end of 1 stage, 2e9 ions (22mA, 2401 kHz RF)
Preamplifiers were not wideband enough and were not used.
Position with transformers, BBB synth algorithm, averaged over 100 bunches
ADC with impedance transformers,
Amplifier110 gain=40dB
nm

## Slide 27: Position measurements results
Transformers, sweep from 736kHz to 2401kHz in 360ms , 2e9 ions
Position with transformers, BBB synth algorithm, averaged over 100 bunches
SUM with transformers, BBB synth algorithm, averaged over 100 bunches
nm
Conclusion:
both transformer and preamp can be used for matching impedance. Preamps are better for low intensities, transformers are more robust.
Libera Hadron provides adequate position measurements for booster beam from injection to extraction. NBA algorithm is preferable for low intensities

## Slide 28: Software
Linux running on Inter connection board.
BMC – baseboard management controller
IPMI – Intelligent Platform Management Interface
MCI – Measurement and control interface

## Slide 29: Software
1. Access to all registers and measurements via remote or local linux console.
2. EPICS server and client
3. TANGO device server mapping available registers, sensors and measurements to TANGO attributes. Configurable via xml file.

## Slide 30: Ongoing tasks
Further measurement tests.
RF frequency distribution: repeater and cables to all Libera devices
Trigger and postmortem signals generation and distribution: cables to all Libera devices
Impedance matching circuits for all pickups
Calibration of 32 measurement devices (4 channels each): impedance matching circuit + Amplifer110 + Libera hbpm
Expected completion date - spring 2019.

## Slide 31: Thank you for your attention!