## Slide 1: GVD optical modules and other mechanics: experience and progress
Alexander Doroshenko, INR RAS, for
VLVNT Conference, 2 Oct. 2018, Dubna

## Slide 2
Deepwater modules - experience and progress
Improvements of the elements
Production facilities progress
Deepwater cables - modernization
New string configuration
Communication and carrier cables production
Deployment equipment – evolution
Mounting winches
Ice works
Bottom cable - double laying
2

## Slide 3
Deepwater modules - experience
Overall design of the optical module and electronic modules of Baikal-GVD was developed mainly to 2015
• Numerous leakages (water infiltration inside module housing) due to bad-quality commercial sealing tape, which was commonly used early
• Large number of defective brass nuts for fixing the modules
We had to do:
- express purchase another sealant (recommended) and perform repeated sealing of all modules of Cluster 2 (two weeks spent in EXP-2017)
extensive repair of Cluster 1 in EXP-2017 (replacement of 30 Oms) and EXP-2018 (replacement of 45 OMs, drying and testing of 15 electronic modules at 7 strings), replacement of all brass nuts by stainless steel ones with a copper coating
EXPEDITION-2017: SERIOUS PROBLEMS
Need purchase serial products only from reliable suppliers (sometimes this leads to long waiting times)
Further refinements were based on the experience of operation of the telescope and mass production of the components

## Slide 4
EXPEDITION 2016: We found ~15 such OMs, exfoliation effects were correlated with low ambient temperatures.
Deepwater modules - experience
exfoliations
heterogeneities
bubbles
PROBLEMS ASSOCIATED WITH OPTICAL COMPOUND (SILGEL 612)
wire bubbles
EXP – 2017, 2018: no exfoliation, but a few bubbles (~2% OMs)
CHANGES to THE PROCESS:
The component mixing ratio A/B shifted from 1.5 to 1.4 for better adhesion
The special clamp fixtures designed for positioning of the PMT in hemisphere
- The regulations introduced for the preparation and mixing of the components
- The environmental conditions control performed
- Treatment with a special primer applied for a permalloy shield (since 2018)

## Slide 5
2016: found a lot of cracks around the holes due to bearing stress from connectors.
Deepwater modules - experience
PROBLEMS CAUSED BY MECHANICAL PROPERTIES OF THE GLASS SPHERES
EXP – 2017, 2018: no cracks, no leakages
Until 2016 we used buoyancies VITROVEX ® 17’’x14mm and made all the holes ourselves.
Later we began to buy such spheres with 2 holes and 2 connectors SubConn™ installed.
Glass cracking tests showed: for the same quality of grinding surfaces around the holes old spheres withstand tightening torque of the connector up to15 Nm without cracking, while the new ones begin to crack at 6 Nm (30% holes cracked at 15 Nm).
adapter washer
connector
Baikal
Nautilus
torque control
crack
- Had to use the adapter washer recommended for the case, with slightly changed design: less thickness, better centering
- Had to find a high-skilled manufacturer to produce them

## Slide 6
Deepwater modules - progress
Reserves of crucial hardware components are forming
Since 2016 we establish successfully cooperation with a small company in Ryazan to produce the wire meshes from permalloy and simultaneously with a large aircraft building plant for their annealing
PREPARING OF MAGNETIC SHIELDS
Chart: Geomagnetic field attenuation of permalloy shields, 30 pcs. <K> = 5.43
| 1 | 1.1000000000000101 | 1.2000000000000099 | 1.30000000000001 | 1.4000000000000099 | 1.50000000000001 | 1.6000000000000101 | 1.7000000000000099 | 1.80000000000001 | 1.9000000000000099 | 2.0000000000000102 | 2.1000000000000099 | 2.2000000000000099 | 2.30000000000001 | 2.4000000000000101 | 2.5000000000000102 | 2.6 | 2.7 | 2.8 | 2.9 | 3 | 3.1 | 3.2 | 3.3 | 3.4 | 3.5 | 3.6 | 3.7 | 3.8 | 3.9 | 4 | 4.0999999999999996 | 4.2 | 4.3 | 4.4000000000000004 | 4.5 | 4.5999999999999996 | 4.7 | 4.8 | 4.9000000000000004 | 5 | 5.0999999999999996 | 5.2 | 5.3 | 5.4 | 5.5 | 5.6 | 5.7 | 5.8 | 5.9 | 6 | 6.1 | 6.2 | 6.3 | 6.4 | 6.5 | 6.6 | 6.7 | 6.8 | 6.9 | 7
N, pcs | 1 | 2 | 3 | 1 | 2 | 8 | 8 | 4 | 1
Now we have in stock
~2000 shield sets ready to annealing ~500 sets ready to use

## Slide 7
Deepwater modules - progress
Reserves of crucial hardware components are forming
2016-2017: The special automated 4-coordinates drilling-grinding machine for the 17” glass hemispheres was designed and built.
HOUSINGS FOR ELECTRONIC MODULES PRODUCTION
Now we have in stock
~ 150 drilled hemisphere
~ 50 housings ready to use
• It drills the hole and forms the plane at any point of the envelope that you need
• The holes have a high quality and accuracy, exceed the established demands
• Manual finishing remains for the grinding plane and the chamfers
• Installation of connectors and two-sided darkening of the glass housings is performed by one trained specialist

## Slide 8
Deepwater modules - progress
Equipment allows to cope with the irregular supply of components
JINR FACILITIES FOR THE OPTICAL MODULES PRODUCTION
- PMT clamp fixtures: 12 sets
- Dozing and mixing system
- Central vacuum, air
- Nitrogen drying system for OM
Test line, 12 OM/day
Assemble line, 12 OM/day
Now we have
~ 450 OMs ready to use

## Slide 9
String structure is significantly optimized
In the sector from the anchor to the String electronic Module (SeM), all deep-water equipment is held on the steel geophysical cable of a small diameter.
For the sector above SeM up to buoy, we use a special carrier cable containing a coaxial pair and supply line.
STRING CONFIGURATION
Deepwater cables - modernization
CHANGES in 2018:
1. SeM moved to the top of the string
+ unification and acceleration of the string assembling
+ simple access to SeM for adjustment
-- longer (heavier-to-handle) some of the patch cables between Central Modules (CM) and SeM
2. Acoustic Modems (AM) connected to CM and SeM by the separate cables instead of “acoustic braid”
+ increasing reliability of acoustic positioning system
+ acceleration of the string assembling
2016-2017
2018
CM
SeM
SeM
CM
CM
CM
CM
CM
acoustic braid

## Slide 10
Deepwater cables - modernization
CABLES PRODUCTION - OM
OM cable - most common cable of Baikal-GVD
Connects OM with CM, is part of the spectrometric channel
Includes coaxial pair, power supply, slow communication line
Has 2 hermetic connectors, 2 hermetic sleeves, length 90 m
Until recently, the whole assortment of linear underwater cable for Baikal-GVD was produced exclusively by PskovGeocabel’ factory.
Further processing to obtaining required hermetic kits was made by our Irkutsk group.
2017: Pilot batch of 16 OM cables fabricated at the plant Expocabel’ (Podolsk)
- successfully tested in Expedition-2018 and installed in Telescope
- Calibration technique agreed with the plant
- Technical Documentation approved
- Facility for the mass production and calibration created at the plant
- Production contract on 350 pcs prepared
2018:
Calibration ill.
Ready-to-use industrial products are created

## Slide 11
Deepwater cables - modernization
CABLES PRODUCTION - CARRIER
String carrier - stainless steel armored geophysical cable
The bottom part of the buoy rope of a length ~700 m
Consist of the core with 7 electrical lines and 2 armoring layers of stainless steel wires
Diameter 9.3 mm, breaking load 43 kN
2017: We placed order for the production at the plant PermGeocabel’ (Perm’)
- The batch produced covered the needs of 2 Clusters. Cluster #3 is assembled on it.
- Tests in Expedition-2018 showed the better resistance against twisting under load in comparison with the previous cable.
2018:
- Design slightly modified for the higher stability to low ambient temperatures
- Plant began work on applying the reliable marking to the cable with a cold gas-dynamic copper sprinkling
Production contract on the batch for 2 Clusters prepared
Ready-to-use industrial products are created

## Slide 12
Deployment resources - evolution
Ice work productivity increasing
MOUNTING WINCHES
Working height 5m and 6m
2-speed transmission
PS ~3x220V, 7 kW
Working load to 20 kN,
Weight ~3500 kg
2018: 6 winches from the basic 10 were equipped with universal reels and sheaves.
Now they can be used with both parts of the carrier cable: thing geophysical steel cable and thick one, with plastic sheath, equipped with hermetic end sleeves.

## Slide 13
ICE WORKS
Deployment resources - evolution
Some steps to avoid work that can be avoided:
- Single-use cable reels and cassettes (realized)
- Carrier cable marking by the plant (in process)
disbanding
disbanding
Ice work productivity increasing
marking
Auxiliary works in the expedition takes up most of the time and human resources
bundling

## Slide 14
BOTTOM CABLE - DOUBLE LAYING TO BE IN 2019
Hybrid optical-electrical carrier cable with plastic sheath
3 single-mode optic fibers in stainless tube (x2)
3+1 electric lines of total section ~16 mm²
Working load to 18kN, breaking load 65 kN
Diameter 19.2 mm, length ~6,5 km
Deployment resources - evolution
Cable was specially designed in 2015 to meet demands of Baikal-GVD project.
It is produced by PskovGeocabel’. One cable serves one Cluster.
2011, 2016, 2017 - Made cabling, one in expedition
2018, status:
Maintenance and repair of the laying equipment performed
Second laying winch and cable transporter prepared
Reserve cutting machine prepared
Two bottom cables produced (at the factory now)
There is a willingness of the team to do this work

## Slide 15
Deployment resources - evolution
possibilities for attracting of manpower are growing up
New technical capabilities
New living modules

## Slide 16
Thank you!
The experience of recent years has allowed to optimize many essential things
We see that in the absence of financial and weather anomalies, the Baikal-GVD team can confidently put into operation 2 Clusters per year
Conclusions
underwater site sonogram, 2017